| CHEMTECH PROJECT NUMBER: O5252 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 15% in the Initial Calibration method (Method 8270-BF103023.M) for 2,4-Dinitrophenol , 4,6-Dinitro-2-methylphenol, Benzo(g,h,i)perylene , these compounds are passing on Linear Regression. The % RSD is greater than 15% in the Initial Calibration method (Method 8270-BF110723.M) for 2,4-Dinitrophenol,4,6-Dinitro-2-methylphenol,these compounds are passing on Linear Regression. The % RSD is greater than 15% in the Initial Calibration method (Method 8270-BM103023.M) for 2-Nitrophenol ,Hexachlorocyclopentadiene ,2-Nitroaniline, 3-Nitroaniline , 2,4-Dinitrophenol,4-Nitrophenol, 2,4-Dinitrotoluene, 4-Nitroaniline ,4,6-Dinitro-2-methylphenol, Dibenzo(a,h)anthracene, these compounds are passing on Linear Regression The Continuous Calibration File ID BM042685.D met the requirements except for 4-Bromophenyl-phenylether, 4-Chlorophenyl-phenylether, Hexachlorobenzene, Hexachlorobutadiene, Hexachlorocyclopentadiene, n-Nitroso-di-n-propylamine,2,4,6-Tribromophenol and Terphenyl-d14 but no positive hits in associated sample therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MSD {O5257-05MSD} with File ID: BF136163.D recoveries met the acceptable requirements except for 3-Nitroaniline[63%], 4-Chloroaniline[24%], Atrazine[145%] and Benzo(g,h,i)perylene[60%] due to matrix interference. The Blank Spike for {PB156921BS} with File ID: BF136178.D met requirements for all samples except for 4-Chloroaniline[59%] this compound did not meet the NJDKQP criteria but met the in-house criteria. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples WASTE was diluted viscous and dirty matrix.

The Form 6 is not included in the data package because the Initial Calibration was performed using 8 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date