CASE NARRATIVE

### LiRo Engineers, Inc.

### Project Name: Walter Gladwin Recreation Center, Bronx, NY

### Project # N/A

### Chemtech Project # P1747

Test Name: PCB

A. Number of Samples and Date of Receipt:

10 Water samples were received on 03/13/2024.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

Anions Group1, CBOD5, Chloride, Dissolved ICP-TAL Metals, Dissolved Mercury, DISSOLVED METALS-TAL, Flash Point, Hexavalent Chromium, Mercury, Metals ICP-TAL, METALS-NYCD, METALS-TAL, Non-Polar Material, NYCDischarge, PCB, Pesticide-TCL, Phenolics, SVOC-NYCD, SVOC-TCL BNA -20, TKN, Total  Nitrogen, TS, TSS, VOC-NYCD and VOC-TCLVOA-10. This data package contains results for PCB.

C. Analytical Techniques:

The analyses were performed on instrument GCECD_O. The front column is ZB-MR1 which is 30 meters, 0.32 mm ID, 0.5 um df, Catalogue # 7HM-G016-17. The rear column is ZB-MR2 which is 30 meters, 0.32 mm ID, 0.25 µm; Catalogue # 7HM-G017-11.The analysis of PCBs was based on method 608.3,8082A and extraction was done based on method 3510.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries met the acceptable criteria except for MW-01 [Tetrachloro-m-xylene(1) - 54%],AS per method one surrogate allowed to fail to meet the criteria per column, No further corrective action was taken.

The Retention Times were acceptable for all samples.

The RPD met criteria .

The Blank Spike met requirements for all samples .

The Blank Spike Duplicate met requirements for all samples .

The Blank analysis did not indicate the presence of lab contamination.

The Initial Calibration met the requirements .

The Continuous Calibration File ID PO102519.D met the requirements except for Aroclor-1260(Peak-05),Decachlorobiphenyl is failing in 2nd column,

The Continuous Calibration File ID PO102531.D met the requirements except for Aroclor-1260(Peak-05),Decachlorobiphenyl is failing in 2nd column,

Above CCAL failing high in 2nd column but it is passing in 1st column therefore no corrective action taken.

E. Additional Comments:

As per method, MS/MSD is required to be performed with the sample analysis. However, Lab did not receive sufficient volume to perform the MS/MSD therefore MS/MSD were not performed for this project. However, Lab has performed LCS/LCSD instead.

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________