CASE NARRATIVE

### LiRo Engineers, Inc.

### Project Name: Walter Gladwin Recreation Center, Bronx, NY

### Project # N/A

### Chemtech Project # P1747

Test Name: SVOC-NYCD

A. Number of Samples and Date of Receipt:

10 Water samples were received on 03/13/2024.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

Anions Group1, CBOD5, Chloride, Dissolved ICP-TAL Metals, Dissolved Mercury, DISSOLVED METALS-TAL, Flash Point, Hexavalent Chromium, Mercury, Metals ICP-TAL, METALS-NYCD, METALS-TAL, Non-Polar Material, NYCDischarge, PCB, Pesticide-TCL, Phenolics, SVOC-NYCD, SVOC-TCL BNA -20, TKN, Total  Nitrogen, TS, TSS, VOC-NYCD and VOC-TCLVOA-10. This data package contains results for SVOC-NYCD.

C. Analytical Techniques:

The samples were analyzed on instrument BNA_F using GC Column DB-UI 8270D which is 20 meters, 0.18 mm ID, 0.36 um dfThe samples were analyzed on instrument BNA_M using GC Column ZB-SemiVolatiles Guardian which is 30 meters, 0.25 mm ID, 0.5 um df, Catalog # 7HG-G027-17-GGAThe analysis of SVOC-NYCD was based on method 625.1 and extraction was done based on method 3510.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries met the acceptable criteria except for

MW-01 [2-Fluorophenol - 42%, Phenol-d6 - 23%],

MW-01RE [2-Fluorophenol - 40% and Phenol-d6 - 22%], All the failure samples in surrogates were reanalyzed to confirm the results as per method and reported in the data.

The Internal Standards Areas met the acceptable requirements.

The Retention Times were acceptable for all samples.

The RPD met criteria .

The Blank Spike met requirements for all samples .

The Blank Spike Duplicate met requirements for all samples .

The Blank analysis did not indicate the presence of lab contamination.

The Initial Calibration met the requirements .

The Continuous Calibration met the requirements .

The Tuning criteria met requirements.

E. Additional Comments:

The Form 6 is not included in the data package because the Initial Calibration was performed using 8 points.

As per method, MS/MSD is required to be performed with sample analysis. However, Lab did not receive sufficient volume to perform the MS/MSD therefore MS/MSD were not performed for this project. However, Lab has performed LCS/LCSD instead.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________