SDG NARRATIVE

### USEPA

### SDG # MYD599

### CASE # 51495

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### CHEMTECH  PROJECT #P2797

### MODIFIED ANALYSIS#3208.0

A. Number of Samples and Date of Receipt

15 Soil and 05 Water samples were delivered to the laboratory intact on 06/08/2024, 06/10/2024

## B. Parameters

# Test requested for Metals CLP Full = Antimony, Arsenic, Barium, Beryllium, Cadmium, Chromium, Cobalt, Copper, Lead, Manganese, Molybdenum, Nickel, Selenium, Silver, Thallium, Vanadium & Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.1oC, 22.9oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1: A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: The laboratory received samples without ice. The coolers had temperatures 24.2 degrees C, 23.2 degrees C, 23.8 degrees C, 24.1 degrees C, and 26.1 degrees C upon arrival. The laboratory would like to know how to proceed.

E. Corrective Action taken for above:

Resolution 1: To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: Per Region 9, Case 51495 is for metals. There are no rinsates in those cooler so they don’t require ice. The laboratory should note the issue in the SDG narrative and proceed with the analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

Inter Element correction factors (IECs) are determined annually and correction factor are applied during ICP-AES analysis.

G. Calculation:

Calculation for ICP-AES Soil Sample:

Conversion of Results from mg/L or ppm to mg/kg (Dry Weight Basis):

Concentration (mg/kg) =     C  x      Vf      x  DF

W x S

Where,

C   = Instrument value in ppm (The average of all replicate exposures)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample MYD4P3  For Arsenic:

If C  = 0.9141476 ppm

Vf = 100 ml

W  = 1.15g

S    = 0.982(98.2/100)

DF = 1

Concentration (mg/kg)   =  0.9141476 x            100       x 1

1.15 x 0.982

=  80.94816 mg/kg

=  81  mg/kg (Reported Result with Signification)

Calculation for ICP-AES Water Sample:

Concentration or Result (μg/L) =     C  x    Vf    x  DF x 1000

Vi

Where,

C  = Instrument value in ppm (The average of all replicate exposures)

Vf = Final digestion volume (mL)

Vi = Initial aliquot amount (mL) (Sample amount taken in prep)

DF = Dilution Factor

Example Calculation For Sample MYD599 For Manganese:

If C    = 0.0020914 ppm

Vf = 50 ml

Vi  = 50 ml

DF = 1

Concentration or Result (μg/L) =     0.0020914 x   50   x 1 x 1000

50

= 2.0914 μg/L

=   2.1 μg/L  (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet

requirements except for Antimony, Selenium, Silver, Thallium. Duplicate sample did meet requirements.

Serial Dilution did meet requirements except for Chromium and Manganese.

Chemical or physical interference effect was suspected and the data for all affected analytes in the sample

received and associated with this serial dilution were flagged.

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer