| CHEMTECH PROJECT NUMBER: P3609 | MATRIX: Water |
| --- | --- |
| METHOD: 8270-Modified/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID BN033465.D met the requirements except for 2,4,6-Tribromophenol, The failure compound not associated with the client parameters list, therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for 915-J-WS-081424 [Terphenyl-d14 - 163%], this compound did not meet the NJDKQP criteria and in-house criteria, and as per method two surrogates are allowed to failed, no corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike met requirements for all samples . The Blank Spike Duplicate met requirements for all samples . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date