| CHEMTECH PROJECT NUMBER: P3847 | MATRIX: Water |
| --- | --- |
| METHOD: 6010D,6020B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 7. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag, Cd, Sn, Sb, Ba, Tl, Pb, U are being analyzed with Non-Collision Cell. Helium gas is used for the Collision Cell analysis.

Internal standard 45Sc(1) and (2) was out Side qc limit for samples P3847-05 in Original so for this sample affected parameters are reported from 5X Dilution.

_______________________________________				__________________________

QA REVIEW								Date