| CHEMTECH PROJECT NUMBER: P4460 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF101824.M) for 2,4-Dinitrophenol, this compound is passing on Linear Regression The Continuous Calibration File ID BF139952.D met the requirements except for 2,4-Dinitrophenol and 4,6-Dinitro-2-methylphenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken. . The Continuous Calibration File ID BF140001.D met the requirements except for 2,4-Dinitrophenol and 4,6-Dinitro-2-methylphenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken. The Continuous Calibration File ID BF140050.D met the requirements except for 2,4-Dinitrophenol and 4,6-Dinitro-2-methylphenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4460-03MS} with File ID: BF140006.D recoveries met the requirements for all compounds except for 4,6-Dinitro-2-methylphenol[148%], 4-Chloroaniline[34%] and Hexachlorocyclopentadiene[164%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. The MSD {P4460-03MSD} with File ID: BF140007.D recoveries met the acceptable requirements except for 4,6-Dinitro-2-methylphenol[157%], 4-Chloroaniline[36%], Di-n-octyl phthalate[138%], these compounds did not meet the NJDKQP criteria but met the in-house criteria, while Hexachlorocyclopentadiene[179%], Atrazine[133%], Dibenz(a,h)anthracene[138%], Indeno(1,2,3-cd)pyrene[138%], these compounds did not meet the NJDKQP criteria and in-house criteria, due to matrix interference no corrective action is required. The Blank Spike for {PB164286BS} with File ID: BF139963.D met requirements for all samples except for 4-Chloroaniline[65%], Hexachlorocyclopentadiene[161%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. The Blank Spike for {PB164369BS} with File ID: BM048237.D met requirements for all samples except for 3-Nitroaniline[60%], 4-Chloroaniline[40%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. while Hexachlorocyclopentadiene[210%], this compound did not meet the NJDKQP criteria and in-house criteria, but associate samples have no positive hit for these compounds therefore no corrective action was taken. The Blank Spike Duplicate for {PB164369BSD} with File ID: BM048233.D met requirements for all samples except for 3,3-Dichlorobenzidine[69%], 3-Nitroaniline[58%], 4-Chloroaniline[41%], these compounds did not meet the NJDKQP criteria but met the in-house criteria, while Hexachlorocyclopentadiene[210%], this compound did not meet the NJDKQP criteria and in-house criteria, but associate samples have no positive hit for these compounds therefore no corrective action was taken . The RPD for {P4460-03MSD} with File ID: BF140007.D met criteria except for Benzo(g,h,i)perylene[38%], Dibenz(a,h)anthracene[37%] and Indeno(1,2,3-cd)pyrene[37%], these compounds did not meet the NJDKQP criteria and in-house criteria due to matrix interference, no corrective action is required. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date