| CHEMTECH PROJECT NUMBER: P4518 | MATRIX: Solid |
| --- | --- |
| METHOD: NJEPH/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for EX-5-TPH-1DL [2-Bromonaphthalene (SURR) - 156%], EX-5-TPH-2 [1-chlorooctadecane (SURR) - 37%], EX-5-TPH-2DL [1-chlorooctadecane (SURR) - 34%], EX-5-TPH-2DL2 [1-chlorooctadecane (SURR) - 38%], EX-5-TPH-2DL2 [2-Bromonaphthalene (SURR) - 141%], EX-5-TPH-4 [2-Flurobiphenyl (SURR) - 144%], EX-5-TPH-4DL [2-Bromonaphthalene (SURR) - 154%] and EX-5-TPH-4DL2 [2-Bromonaphthalene (SURR) - 164%] due to high concentration, samples required dilution therefore samples were reanalyzed with dilution and reported. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4518-05MS} with File ID: FC067546.D recoveries met the requirements for all compounds except for Aliphatic C9-C12[145%], Aliphatic C12-C16[921%], Aliphatic C16-C21[762%], Aliphatic C21-C28[377%] and Aliphatic C28-C40[321%] due to sample matrix interference. The MS {P4518-05MS} with File ID: FD048607.D recoveries met the requirements for all compounds except for Aromatic C12-C16[223%], Aromatic C16-C21[281%] due to sample matrix interference. The MSD {P4518-05MSD} with File ID: FC067547.D recoveries met the acceptable requirements except for Aliphatic C9-C12[140%], Aliphatic C12-C16[856%], Aliphatic C16-C21[748%], Aliphatic C21-C28[350%] and Aliphatic C28-C40[307%] due to sample matrix interference. The MSD {P4518-05MSD} with File ID: FD048608.D recoveries met the acceptable requirements except for Aromatic C12-C16[180%], Aromatic C16-C21[277%] due to sample matrix interference. The Blank Spike met requirements for all samples . The Blank Spike Duplicate met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date