| CHEMTECH PROJECT NUMBER: P4578 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration method (Method 8270-BE102824.M) for 2,4-Dinitrophenol, this compound is passing on Linear Regression. The % RSD is greater than 20% in the Initial Calibration (8270-BF101824.M) for 2,4-Dinitrophenol, this compound is passing on Linear Regression. The Continuous Calibration File ID BF140100.D met the requirements except for 2,4-Dinitrophenol and 4,6-Dinitro-2-methylphenol are failing high but no positive hit in associate sample therefore no corrective action taken. The Continuous Calibration File ID BF140113.D met the requirements except for 2,4-Dinitrophenol,2-Nitrophenol,4,6-Dinitro-2-methylphenol,Bis(2-ethylhexyl)phthalate and Di-n-octyl phthalate are failing high but no positive hit in associate sample therefore no corrective action taken. The Continuous Calibration File ID BF140152.D met the requirements except for 2,4-Dinitrophenol,2,4-Dinitrotoluene,2-Nitrophenol,4,6-Dinitro-2-methylphenol,4-Nitroaniline,Bis(2-ethylhexyl)phthalate,Butylbenzylphthalate and Di-n-octyl phthalate are failing high but no positive hit in associate sample therefore no corrective action taken. The Continuous Calibration File ID BF140176.D met the requirements except for 2,4-Dinitrophenol,2,4-Dinitrotoluene,2-Nitrophenol,4,6-Dinitro-2-methylphenol,4-Nitroaniline,Bis(2-ethylhexyl)phthalate,Butylbenzylphthalate and Di-n-octyl phthalate are failing high but no positive hit in associate sample therefore no corrective action taken. The Continuous Calibration File ID BF140200.D met the requirements except for 2,4-Dinitrophenol,2,4-Dinitrotoluene,2-Nitrophenol,4,6-Dinitro-2-methylphenol,4-Nitroaniline,4-Nitrophenol,Bis(2-ethylhexyl)phthalate,Butylbenzylphthalate and Di-n-octyl phthalate are failing high but no positive hit in associate sample therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for WB-305-SW [2-Fluorobiphenyl - 19%, 2-Fluorophenol - 9%, Nitrobenzene-d5 - 17%, Phenol-d6 - 8%, Terphenyl-d14 - 23%], WB-305-SWRE [2-Fluorobiphenyl - 17%, 2-Fluorophenol - 9%, Nitrobenzene-d5 - 17%, Phenol-d6 - 9%, Terphenyl-d14 - 24%], WB-305-SW-1RE [2-Fluorobiphenyl - 23%, 2-Fluorophenol - 14%, Nitrobenzene-d5 - 24%, Phenol-d6 - 12%, Terphenyl-d14 - 27%] these surrogates did not meet the NJDKQP criteria and in-house criteria While WB-305-SW [2,4,6-Tribromophenol - 20%], WB-305-SWRE [2,4,6-Tribromophenol - 23%], WB-305-SW-1 [Nitrobenzene-d5 - 43%, 2-Fluorobiphenyl - 46%], WB-305-SW-1RE [2,4,6-Tribromophenol - 28%], these surrogates did not meet the in-house criteria but met the NJDKQP criteria, all samples were reanalyzed to confirm the failure and reported. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4561-05MS} with File ID: BF140105.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[58%], 3-Nitroaniline[50%], 4,6-Dinitro-2-methylphenol[142%], 4-Chloroaniline[23%], Benzo(g,h,i)perylene[69%] and Hexachlorocyclopentadiene[170%] these compounds did not meet the NJDKQP criteria but met the in-house criteria . The MSD {P4561-05MSD} with File ID: BF140106.D recoveries met the acceptable requirements except for 3,3-Dichlorobenzidine[61%], 3-Nitroaniline[49%], 4,6-Dinitro-2-methylphenol[142%], 4-Chloroaniline[23%], Benzo(g,h,i)perylene[68%] and Hexachlorocyclopentadiene[161%] these compounds did not meet the NJDKQP criteria but met the in-house criteria . The Blank Spike for {PB164533BS} with File ID: BE101427.D met requirements for all samples except for 4-Chloroaniline[66%] this compound did not meet the NJDKQP criteria but met the in-house criteria while Hexachlorocyclopentadiene[200%] this compound did not meet the NJDKQP criteria and in-house criteria, are failing high but no positive hit in associate sample therefore no corrective action taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date