| CHEMTECH PROJECT NUMBER: P4578 | MATRIX: TCLP |
| --- | --- |
| METHOD: 8151A/3510/1311 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for WB-305-BOT [2,4-DCAA(1) - 68%, 2,4-DCAA(2) - 56%], WB-305-BOTMS [2,4-DCAA(2) - 58%], WB-305-BOTMSD [2,4-DCAA(2) - 60%], WB-305-BOT-1 [2 and4-DCAA(2) - 61%], PB164393TB[2 and4-DCAA(2) - 58%],these compounds did not meet the NJDKQP criteria but met the in-house criteria . |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4578-06MS} with File ID: PS028291.D recoveries met the requirements for all compounds except for 2,4,5-TP(Silvex)[144%]this compound did not meet the NJDKQP criteria and in-house criteria due to matrix interference. The MSD {P4578-06MSD} with File ID: PS028292.D recoveries met the acceptable requirements except for 2,4,5-TP(Silvex)[143%],this compound did not meet the NJDKQP criteria and in-house criteria due to matrix interference. The Blank Spike met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. The Holding Times were met for all analysis. |  |  |  |

ADDITIONAL COMMENTS:

_______________________________________				__________________________

QA REVIEW								Date