| CHEMTECH PROJECT NUMBER: P4599 | MATRIX: Solid |
| --- | --- |
| METHOD: NJEPH/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except For Aromatic EX-1-TPH-1 [2-Flurobiphenyl (SURR) - 154%], EX-1-TPH-1DL [2-Bromonaphthalene (SURR) - 146%, 2-Flurobiphenyl (SURR) - 141%], EX-1-TPH-1DL2 [2-Bromonaphthalene (SURR) - 219%, 2-Flurobiphenyl (SURR) - 212%], EX-1-TPH-2 [2-Bromonaphthalene (SURR) - 155%, 2-Flurobiphenyl (SURR) - 174%], EX-1-TPH-2DL [2-Bromonaphthalene (SURR) - 168%, 2-Flurobiphenyl (SURR) - 164%], EX-1-TPH-2DL2 [2-Bromonaphthalene (SURR) - 215%, 2-Flurobiphenyl (SURR) - 214%], EX-1-TPH-3 [2-Bromonaphthalene (SURR) - 258%, 2-Flurobiphenyl (SURR) - 170%], EX-1-TPH-3DL [2-Bromonaphthalene (SURR) - 242%, 2-Flurobiphenyl (SURR) - 242%], EX-1-TPH-3DL2 [2-Bromonaphthalene (SURR) - 324%, 2-Flurobiphenyl (SURR) - 288%], EX-1-TPH-4 [2-Bromonaphthalene (SURR) - 178%, 2-Flurobiphenyl (SURR) - 205%], EX-1-TPH-4DL [2-Bromonaphthalene (SURR) - 198%, 2-Flurobiphenyl (SURR) - 193%], EX-1-TPH-4DL2 [2-Bromonaphthalene (SURR) - 235%, 2-Flurobiphenyl (SURR) - 226%], EX-1-TPH-5 [2-Bromonaphthalene (SURR) - 176%, 2-Flurobiphenyl (SURR) - 150%], EX-1-TPH-5DL [2-Bromonaphthalene (SURR) - 206%, 2-Flurobiphenyl (SURR) - 212%], EX-1-TPH-5DL2 [2-Bromonaphthalene (SURR) - 259%, 2-Flurobiphenyl (SURR) - 259%], EX-1-TPH-6 [2-Bromonaphthalene (SURR) - 155%], EX-1-TPH-6DL [2-Bromonaphthalene (SURR) - 164%, 2-Flurobiphenyl (SURR) - 146%], EX-1-TPH-8 [2-Bromonaphthalene (SURR) - 215%, 2-Flurobiphenyl (SURR) - 203%], EX-1-TPH-8DL [2-Bromonaphthalene (SURR) - 248%, 2-Flurobiphenyl (SURR) - 206%], EX-1-TPH-9 [2-Bromonaphthalene (SURR) - 218%, 2-Flurobiphenyl (SURR) - 197%], EX-1-TPH-9DL [2-Bromonaphthalene (SURR) - 202%, 2-Flurobiphenyl (SURR) - 188%, ortho-Terphenyl (SURR) - 141%], EX-1-TPH-9DL2 [2-Bromonaphthalene (SURR) - 217%, 2-Flurobiphenyl (SURR) - 200% and ortho-Terphenyl (SURR) - 153%],due to high concentration while, samples required dilution therefore samples were reanalyzed with dilution and reported. For Aliphatic EX-1-TPH-1DL [1-chlorooctadecane (SURR) - 0%], EX-1-TPH-2DL [1-chlorooctadecane (SURR) - 0%], EX-1-TPH-3DL2 [1-chlorooctadecane (SURR) - 0%],EX-1-TPH-4DL2 [1-chlorooctadecane (SURR) - 0%],EX-1-TPH-5DL2 [1-chlorooctadecane (SURR) - 0%], EX-1-TPH-8DL2 [1-chlorooctadecane (SURR) - 0%], EX-1-TPH-9DL2 [1-chlorooctadecane (SURR) - 0%],Surrogates were diluted out due to the high dilution. No further corrective action was taken |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS recoveries met the requirements for all compounds . The MSD recoveries met the acceptable requirements . The Blank Spike met requirements for all samples . The Blank Spike Duplicate met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples EX-1-TPH-1, EX-1-TPH-2, EX-1-TPH-3, EX-1-TPH-4, EX-1-TPH-5, EX-1-TPH-6, EX-1-TPH-8 and EX-1-TPH-9, EX-1-TPH-1DL, EX-1-TPH-2DL, EX-1-TPH-3DL, EX-1-TPH-4DL, EX-1-TPH-5DL, EX-1-TPH-8DL, EX-1-TPH-9DL, were diluted due to high concentrations.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date