SDG NARRATIVE

### USEPA

### SDG # MC0VG5

### CASE # 51810

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # P4654

A. Number of Samples and Date of Receipt

19 Soil samples were delivered to the laboratory intact on 10/31/2024, 11/02/2024

B. Parameters

# Test requested for Mercury.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.4oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1 : A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

E. Corrective Action taken for above:

Resolution 1 : To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

G. Calculation:

Calculation for Hg Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument response in μg/L from the calibration curve.

Vf = Final prepared (absorbing solution) volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample MC0VG5:

If C    = 1.0674 ppb

Vf = 100 mL

W  =  0.60g

S    = 0.857(85.7/100)

DF = 1

Concentration (mg/kg) =     1.0674 x         100          x 1 / 1000

0.60 x 0.857

=  0.207584 mg/kg

=   0.21 mg/kg (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Blank analyses did not indicate any presence of contamination. Spike sample did meet requirements. Duplicate sample did meet requirements.

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer