| CHEMTECH PROJECT NUMBER: P4663 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (Method 8270-BE102824.M) for 2,4-Dinitrophenol, this compound is passing on Linear Regression. The % RSD is greater than 20% in the Initial Calibration ( Method 8270-BF101824.M) for 2,4-Dinitrophenol, this compound is passing on Linear Regression. The Continuous Calibration File ID BF140200.D met the requirements except for 2,4-Dinitrophenol, 2,4-Dinitrotoluene, 2-Nitrophenol, 4,6-Dinitro-2-methylphenol, 4-Nitroaniline, 4-Nitrophenol, Bis(2-ethylhexyl)phthalate, Butylbenzylphthalate and Di-n-octyl phthalate but no positive hits in associated samples therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4663-08MS} with File ID: BF140209.D recoveries met the requirements for all compounds except for 4,6-Dinitro-2-methylphenol[139%], 4-Chloroaniline[43%], bis(2-Ethylhexyl)phthalate[167%], Di-n-octyl phthalate[133%], Hexachlorocyclopentadiene[164%] these compounds did not meet the NJDKQP criteria but met the in-house criteria and Butylbenzylphthalate[150%] this compound did not meet the NJDKQP criteria and in-house criteria due to matrix interference. The MSD {P4663-08MSD} with File ID: BF140210.D recoveries met the acceptable requirements except for 4,6-Dinitro-2-methylphenol[150%], 4-Chloroaniline[46%], Di-n-octyl phthalate[144%], Hexachlorocyclopentadiene[169%] these compounds did not meet the NJDKQP criteria but met the in-house criteria and Butylbenzylphthalate[167%], bis(2-Ethylhexyl)phthalate[183%] these compounds did not meet the NJDKQP criteria and in-house criteria due to matrix interference. The Blank Spike for {PB164593BS} with File ID: BE101455.D met requirements for all samples except for 3-Nitroaniline[65%] this compound did not meet the NJDKQP criteria but met the in-house criteria and Hexachlorocyclopentadiene[215%] this compound did not meet the NJDKQP criteria and in-house criteria but no positive hits in associated samples therefore no corrective action taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

For sample # TENNANT-PAD-2-3-STOCKPILE some compounds below Method detection limits, therefore it is not reported as Hit in Form-1.

The Form 6 is not included in the data package because the Initial Calibration was performed using 8 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date