| CHEMTECH PROJECT NUMBER: P4718 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF110524.M) for 2,4-Dinitrophenol, this compound is passing on Linear Regression. The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4718-02MS} with File ID: BF140272.D recoveries met the requirements for all compounds except for Atrazine[143%], Hexachlorocyclopentadiene[198%] these compounds did not meet the NJDKQP criteria and in-house criteria, due to matrix interference. The MSD {P4718-02MSD} with File ID: BF140273.D recoveries met the acceptable requirements except for Atrazine[138%], Hexachlorocyclopentadiene[188%] due to matrix interference. The Blank Spike for {PB164702BS} with File ID: BF140263.D met requirements for all samples except for 3-Nitroaniline[56%], 4-Chloroaniline[43%], Benzaldehyde[19%] these compounds did not meet the NJDKQP criteria but met the in-house criteria , while Hexachlorocyclopentadiene [200%] this compound did not meet the NJDKQP criteria and in-house criteria, are failing high but no positive hit in associate sample therefore no corrective action taken. The Blank Spike for {PB164846BS} with File ID: BF140446.D met requirements for all samples except for 3-Nitroaniline[67%], 4-Chloroaniline[65%] these compounds did not meet the NJDKQP criteria but met the in-house criteria while Hexachlorocyclopentadiene[170%] this compound did not meet the NJDKQP criteria and in-house criteria, failing high but no positive hit in associate sample therefore no corrective action taken and Benzaldehyde [0%] this compound did not meet the NJDKQP criteria and in-house criteria, failing low but associate CCAL passing therefore no corrective action taken. The Blank Spike Duplicate for {PB164846BSD} with File ID: BF140447.D met requirements for all samples except for 3,3-Dichlorobenzidine[62%], 3-Nitroaniline[57%], 4-Chloroaniline[45%], Benzaldehyde[14%] these compounds did not meet the NJDKQP criteria but met the in-house criteria while Hexachlorocyclopentadiene[180%] this compound did not meet the NJDKQP criteria and in-house criteria, failing high but no positive hit in associate sample therefore no corrective action taken. The RPD for {PB164846BSD} with File ID: BF140447.D met criteria except for 3,3-Dichlorobenzidine[22%], 4-Chloroaniline[37%], Atrazine[33%] and Benzaldehyde[200%] due to difference in results of BS and BSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date