CASE NARRATIVE

### CHA Companies, Inc.

### Project Name: HRPC

### Project # N/A

### Chemtech Project # P4779

Test Name: VOC-TCLVOA-10

A. Number of Samples and Date of Receipt:

1 Solid sample was received on 11/08/2024.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

PCB, SVOC-TCL BNA -20 and VOC-TCLVOA-10. This data package contains results for VOC-TCLVOA-10.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_Y were done using GC column Rxi-624Sil MS, which is 30 meters, 0.25 mm id, 1.4 um df, Restek Cat. #13868. The Trap was supplied by Supelco, VOCARB 3000, ATOMAX XYZ Concentrator.The analysis of VOC-TCLVOA-10 was based on method 8260D.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries met the acceptable criteria except for SOIL-01 [1,2-Dichloroethane-d4 - 170%], SOIL-01RE [1 and2-Dichloroethane-d4 - 170%]. Sample reanalyzed to confirm results, Original and Reanalysis both are reported.

The Internal Standards Areas met the acceptable requirements except for SOIL-01 and SOIL-01RE. Sample reanalyzed to confirm results, Original and reanalysis both are reported.

The Retention Times were acceptable for all samples.

The RPD for {VY1111SBSD01} with File ID: VY020242.D met criteria except for Acetone[21%] due to difference in BS and BSD concentrations.

The Blank Spike met requirements for all samples .

The Blank Spike Duplicate met requirements for all samples .

The Blank analysis did not indicate the presence of lab contamination.

The Initial Calibration met the requirements .

The Continuous Calibration File ID VY020262.D met the requirements except for 2-Butanone,2-Hexanone,4-Methyl-2-Pentanone and Acetone . But associated sample has not positive hit for these compounds therefore no corrective action was taken.

The Tuning criteria met requirements.

E. Additional Comments:

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________