SDG NARRATIVE

### USEPA

### SDG # GCP96

### CASE # 51874

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # P4787

A. Number of Samples and Date of Receipt

06 Water  samples were delivered to the laboratory intact on 11/08/2024

## B. Parameter

# Test requested for  Metals CLP MS-CLP12= Aluminum,  Arsenic, Cobalt, Copper, Iron, Lead,  Manganese, Vanadium, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.1oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue: A Laboratory QC sample was not listed on the COC. The laboratory would like to use sample GCP96 for Laboratory QC. The laboratory has confirmed that the sample is not a PE or blank.

E. Corrective Action taken for above:

Resolution: Per SFAM01.1 Exhibit A, Section 5.5.4.1, the laboratory will note the issue in the SDG

Narrative and proceed with the analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

G. Calculation:

Calculation for ICP-MS Water Sample:

Concentration or Result (μg/L) =     C  x    Vf    x  DF

Vi

Where,

C    = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

Vi = Initial aliquot amount (mL) (Sample amount taken in prep)

DF = Dilution Factor

Example Calculation For Sample GCP96 For Arsenic:

If C    = 0.21 ppb

Vf = 50 ml

Vi  = 50 ml

DF = 1

Concentration or Result (μg/L) =     0.21  x   50   x 1

50

=   0.21 μg/L

=   0.21  μg/L  (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements. Duplicate sample did meet requirements except for Manganese. Serial Dilution did meet requirements.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Aluminum | 45Sc |
| Arsenic | 89Y |
| Cobalt | 45Sc |
| Copper | 45Sc |
| Iron | 45Sc |
| Lead | 209Bi |
| Manganese | 45Sc |
| Vanadium | 45Sc |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer