| CHEMTECH PROJECT NUMBER: P4822 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF112124.M) for Hexachlorocyclopentadiene, 2,4-Dinitrophenol, these compounds are passing on Linear Regression The Continuous Calibration File ID BF140501.D met the requirements except for Benzaldehyde . But associated samples have not positive hit for this compound therefor no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4821-01MS} with File ID: BF140432.D recoveries met the requirements for all compounds except for 3-Nitroaniline[61%], 4-Chloroaniline[27%], Benzaldehyde[12%] and Benzo(g,h,i)perylene[60%] . These compounds did not meet the NJDKQP criteria but met the in-house criteria . The MSD {P4821-01MSD} with File ID: BF140433.D recoveries met the acceptable requirements except for 3,3-Dichlorobenzidine[65%], 3-Nitroaniline[50%], 4-Chloroaniline[18%], Benzaldehyde[12%], Benzo(g,h,i)perylene[58%], Dibenz(a,h)anthracene[69%] and Indeno(1,2,3-cd)pyrene[68%] . These compounds did not meet the NJDKQP criteria but met the in-house criteria . The Blank Spike for {PB164940BS} with File ID: BF140449.D met requirements for all samples except for 3,3-Dichlorobenzidine[55%], 3-Nitroaniline[51%], 4-Chloroaniline[38%] These compounds did not meet the NJDKQP criteria but met the in-house criteria and for Hexachlorocyclopentadiene[173%] but associated samples have not positive hit for these compounds therefore no corrective action was taken. The RPD for {P4821-01MSD} with File ID: BF140433.D met criteria except for 4-Chloroaniline[40%] due to difference in MS and MSD concentrations. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date