| CHEMTECH PROJECT NUMBER: P4839 | MATRIX: Solid |
| --- | --- |
| METHOD: 8015D/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for EX-9-TPH-14 [Tetracosane-d50 - 0%], EX-4-TPH-5 [Tetracosane-d50 - 0%]. Surrogates were diluted out due to the high dilution. No further corrective action was taken. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P4839-22MS} with File ID: FF015086.D recoveries met the requirements for all compounds except for DRO[43.8%] due to matrix interference . The MSD {P4839-22MSD} with File ID: FF015087.D recoveries met the acceptable requirements except for DRO[47.7%] due to matrix interference. The Blank Spike met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples EX-9-TPH-9, EX-9-TPH-10, EX-9-TPH-11, EX-9-TPH-13, EX-9-TPH-14, EX-9-TPH-15, EX-9-TPH-17, EX-9-TPH-19, EX-9-TPH-20, EX-9-TPH-21, EX-10-TPH-2, EX-10-TPH-3 and EX-4-TPH-5 were diluted due to bad matrices.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date