SDG NARRATIVE

### USEPA

### SDG # ME29F6

### CASE # 51759

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID #P4977

MODIFIED ANALYSIS# 3227.1

A. Number of Samples and Date of Receipt

07 Soil samples were delivered to the laboratory intact on 09/23/2024

## B. Parameters

# Test requested for Metals CLP12 = Aluminum, Calcium, Iron, Magnesium, Potassium, Sodium.

# Test requested for Metals CLP MS FULL= Antimony, Arsenic, Barium, Beryllium, Cadmium, Chromium, Cobalt, Copper, Lead, Manganese, Nickel, Selenium, Silver, Thallium,  Titanium, Vanadium, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.1oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1: A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: The COCs are missing the relinquished by information.

Issue 3: Hg analysis is not scheduled for this Case but is listed on the COC.

Issue 4: The sample jars for soil samples ME29E9, ME29F1 and ME29G2 were broken inside their Ziplock bags, and all sample volume spilled into the bags. The ice water that submerged the soil samples in the bag could result in crosscontamination.

E. Corrective Action taken for above:

Resolution 1: To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory

will  note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: Per Region 5, a revised COC has been provided with the relinquished by information. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 3: Per Region 5, the laboratory will disregard Hg analysis and proceed with the analysis of the samples as scheduled. Please note the issue in the SDG Narrative.

Resolution 4: Per Region 5, the laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples with the risk of cross‐contamination.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

Inter Element correction factors (IECs) are determined annually and correction factor are applied during ICP-AES analysis.

G. Calculation:

Calculation for ICP-AES Soil Sample:

Conversion of Results from mg/L or ppm to mg/kg (Dry Weight Basis):

Concentration (mg/kg) =     C  x      Vf      x  DF

W x S

Where,

C   = Instrument value in ppm (The average of all replicate exposures)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME29F6 For Aluminum:

If C    = 17.85760 ppm

Vf = 100 ml

W  = 1.18g

S    = 0.90(90/100)

DF = 1

Concentration (mg/kg)   =    17.85760 x           100          x 1

1.18 x 0.90

=   1681.50 mg/kg

=  1700   mg/kg (Reported Result with Signification)

Calculation for ICP-MS Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME29F6 For Arsenic:

If C  =   6.75 ppb

Vf = 500 ml

W  = 1.08 g

S    = 0.90(90/100)

DF = 1

Concentration (mg/kg) =     6.75 x        500           x 1 / 1000

1.08 x 0.90

=   3.4722 mg/kg

=    3.5 mg/kg (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements except For Lead. Duplicate sample did meet requirements except for Arsenic, Barium, Chromium, Cobalt, Copper, Lead, Manganese, Nickel, Vanadium, Zinc. Serial Dilution did meet requirements.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag, Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Antimony | 159Tb |
| Arsenic | 89Y |
| Barium | 159Tb |
| Beryllium | 6Li |
| Cadmium | 159Tb |
| Chromium | 45Sc |
| Cobalt | 45Sc |
| Copper | 45Sc |
| Lead | 209Bi |
| Manganese | 45Sc |
| Nickel | 45Sc |
| Selenium | 89Y |
| Silver | 159Tb |
| Thallium | 209Bi |
| Titanium | 45sc |
| Vanadium | 45Sc |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer