| CHEMTECH PROJECT NUMBER: P5002 | MATRIX: Solid |
| --- | --- |
| METHOD: NJEPH/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for EX-7-TPH-1 [2-Bromonaphthalene (SURR) - 140.94%], EX-7-TPH-1DL [2-Bromonaphthalene (SURR) - 153.8%], EX-7-TPH-1DL2 [2-Bromonaphthalene (SURR) - 183.6%], Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P5002-01MS} with File ID: FC067876.D recoveries met the requirements for all compounds except for Aliphatic C12-C16[1305%], Aliphatic C16-C21[823%], Aliphatic C21-C28[245%] and Aliphatic C28-C40[140%] . due to sample is very bad matrix and original sample required dilution. The MS {P5002-01MS} with File ID: FD048827.D recoveries met the requirements for all compounds except for Aromatic C16-C21[227%], Aromatic C21-C36[185%] . due to sample is very bad matrix and original sample required dilution. The MSD {P5002-01MSD} with File ID: FC067877.D recoveries met the acceptable requirements except for Aliphatic C12-C16[1126%], Aliphatic C16-C21[695%] and Aliphatic C21-C28[194%] due to sample is very bad matrix and original sample required dilution. The MSD {P5002-01MSD} with File ID: FD048828.D recoveries met the acceptable requirements except for Aromatic C12-C16[167%], Aromatic C16-C21[328%] and Aromatic C21-C36[204%] due to sample is very bad matrix and original sample required dilution. The Blank Spike met requirements for all samples . The Blank Spike Duplicate met requirements for all samples . The RPD for {P5002-01MSD} with File ID: FD048828.D met criteria except for Aromatic C12-C16[42%], Aromatic C16-C21[36%] due to difference in MS and MSD concentrations. |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Sample EX-7-TPH-1,EX-7-TPH-1DL were diluted due to high concentration for Aliphatic compounds.

Sample EX-7-TPH-1,EX-7-TPH-1DL were diluted due to high concentration for Aromatic compounds.

The temperature of the samples at the time of receipt was 14.2°C.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date