| CHEMTECH PROJECT NUMBER: P5112 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P5112-01MS} with File ID: BF140780.D recoveries met the requirements for all compounds except for Fluoranthene[11%], Phenanthrene[44%], due to matrix interference, therefore no corrective is required. The MSD {P5112-01MSD} with File ID: BF140781.D recoveries met the acceptable requirements except for Fluoranthene[33%], Phenanthrene[50%], due to matrix interference, therefore no corrective is required. The Blank Spike met requirements for all samples . The RPD for {P5112-01MSD} with File ID: BF140781.D met criteria except for Fluoranthene[100%], due to matrix interference, therefore no corrective is required. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for 10TH-ST-SOIL, 10TH-ST-SOILDL reanalyzed with the required dilution and both run were reported in Hard Copy. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Sample 10TH-ST-SOIL was diluted due to high concentration.

As per special requirement for this project form-1 and Hit Summary are reported in mg/kg.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date