SDG NARRATIVE

### USEPA

### SDG # MJNBZ1

### CASE # 51918

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # P5250

### MODIFIED ANALYSIS #3157.0, 3106.0

A. Number of Samples and Date of Receipt

06 Soil samples was delivered to the laboratory intact on 12/11/2024, 12/14/2024.

B. Parameters

# Test requested for Metals CLP MS FULL = Antimony, Arsenic, Barium, Beryllium, Cadmium, Chromium, Cobalt, Copper, Lead, Manganese, Molybdenum, Nickel, Selenium, Silver, Thallium, Vanadium, Zinc & Mercury.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.0oC, 2.1oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1 : A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: There is no sample designated on the COC for laboratory QC for ICP-MS Metals and Hg analyses in SDG MJNBZ1. The laboratory has selected sample MJNBZ2 for laboratory QC for ICP-MS Metals and Hg, and confirms that the sample is not a PT, blank, or rinsate sample.

Issue 3: According to the COC, samples shipped on 12/09/2024, but the collection date for samples MJNBZ3 and MJNBZ4 is 12/11/2024, indicating these samples were collected two days after the ship date. Please advise on how the laboratory may proceed.

E. Corrective Action taken for above:

Resolution 1 : To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: Per SFAM01.1 Exhibit A, Section 5.5.4.1., the laboratory should note the issue in the SDG Narrative and proceed with analysis of the samples.

Resolution 3: Per Region 10, the laboratory should proceed with a sample collection date of 12/06/2024 for samples MJNBZ3 and MJNBZ4. 12/11/2024 was the date that the samples were originally planned to be collected, but these samples were actually collected on 12/06/2024. Please note the issue in the SDG Narrative and proceed with analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

G. Calculation:

Calculation for ICP-MS Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample MJNBZ1 For Arsenic:

If C  =  9.24 ppb

Vf = 500 ml

W  = 1.16 g

S    = 0.894 (89.4/100)

DF = 1

Concentration (mg/kg) =  9.24   x        500          x 1 / 1000

1.16 x 0.894

=  4.4549 mg/kg

=  4.5 mg/kg (Reported Result with Signification)

Calculation for Hg Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument response in μg/L from the calibration curve.

Vf = Final prepared (absorbing solution) volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample MJNBZ1:

If C    =0.4461 ppb

Vf = 100 mL

W  =  0.60g

S    = 0.894(89.4/100)

DF = 1

Concentration (mg/kg) =     0.4461 x         100          x 1 / 1000

0.60 x 0.894

=  0.083165  mg/kg

=   0.083 mg/kg (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. MS Spike sample did meet requirements. Duplicate sample did meet requirements. Serial Dilution did meet requirements.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Antimony | 159Tb |
| Arsenic | 89Y |
| Barium | 159Tb |
| Beryllium | 6Li |
| Cadmium | 159Tb |
| Chromium | 45Sc |
| Cobalt | 45Sc |
| Copper | 45Sc |
| Lead | 209Bi |
| Manganese | 45Sc |
| Molybdenum | 89Y |
| Nickel | 45Sc |
| Selenium | 89Y |
| Silver | 159Tb |
| Thallium | 209Bi |
| Vanadium | 45Sc |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer