SDG NARRATIVE

### USEPA

### SDG # MJNL92

### CASE # 51821

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # P5256

A. Number of Samples and Date of Receipt

01 Soil sample was delivered to the laboratory intact on 12/11/2024

B. Parameters

# Test requested for Metals CLP4 MS = Arsenic, Copper, Lead, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.0oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1 : A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: The attached COC indicates the laboratory should use CLP sample numbers MJNKK5, MJNKN6, MJNKQ9, MJNKW2, MJNL72, and MJNL92 for laboratory QC, but the laboratory only requires one sample for QC per shipment. The laboratory would like to proceed with performing laboratory QC on CLP sample numbers MJNKK5, MJNL72, and MJNL92 and not use the remaining designated samples for QC. Please advise on how the laboratory may proceed.

E. Corrective Action taken for above:

Resolution 1 : To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory

will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: Per Region 10, it is acceptable for the laboratory to proceed with performing laboratory QC on samples MJNKK5, MJNL72, and MJNL92 and the scheduled analyses on the remaining samples. Please note the issue in the SDG Narrative and proceed with analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

G. Calculation:

Calculation for ICP-MS Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample MJNL92 For Arsenic:

If C  =  419.14 ppb

Vf = 500 ml

W  = 1.24 g

S    = 0.457(45.7/100)

DF = 1

Concentration (mg/kg)  =    419.14 x        500          x 1 / 1000

1.24 x 0.457

=  369.8207 mg/kg

=  370 mg/kg (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements. Duplicate sample did meet requirements. Serial Dilution did meet requirements.

Some samples have % solids results less than 50% but more than 30%.  Please see below table for detail. Laboratory has processed these samples according to the SFAM01.1 SOW, Exhibit D, sections 10.1.1.8.

| EPA Sample ID | % Solid |
| --- | --- |
| MJNL92 | 45.7 |
| MJNL92D | 45.7 |
| MJNL92S | 45.7 |

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Arsenic | 89Y |
| Copper | 45Sc |
| Lead | 209Bi |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer