SDG NARRATIVE

### USEPA

### SDG # ME28W7

### CASE # 51847

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # P5259

A. Number of Samples and Date of Receipt

19 Soil samples were delivered to the laboratory intact on 12/11/2024, 12/12/2024.

B. Parameters

Test requested for Metals CLP12 = Aluminum, Calcium, Iron, Magnesium, Potassium, Sodium & Mercury, Cyanide.

Test requested for Metals CLP  MS FULL = Antimony, Arsenic, Barium, Beryllium, Cadmium, Chromium, Cobalt, Copper, Lead,  Manganese, Nickel,  Selenium, Silver,  Thallium, Vanadium, Zinc

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.2oC, 2.0oC, 2.4oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1: A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

E. Corrective Action taken for above:

Resolution 1: To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

Inter Element correction factors (IECs) are determined annually and correction factor are applied during

ICP-AES analysis.

G. Calculation:

Calculation for ICP-AES Soil Sample:

Conversion of Results from mg/L or ppm to mg/kg (Dry Weight Basis):

Concentration (mg/kg) =     C  x      Vf      x  DF

W x S

Where,

C   = Instrument value in ppm (The average of all replicate exposures)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28W7 For Aluminum :

If C    = 93.17755 ppm

Vf = 100 ml

W  = 1.12g

S    = 0.792 (79.2/100)

DF = 1

Concentration (mg/kg)   =    93.17755  x        100          x 1

1.12  x 0.792

=      10504.32  mg/kg

=     11000 mg/kg (Reported Result with Signification)

Calculation for ICP-MS Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28W7 For Arsenic :

If C  =  12.73 ppb

Vf = 500 ml

W  = 1.02 g

S    =0.792 (79.2/100)

DF = 1

Concentration (mg/kg) =     12.73 x        500              x 1 / 1000

1.02 x 0.792

=    7.8790 mg/kg

=   7.9 mg/kg (Reported Result with Signification)

Calculation for Hg Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument response in μg/L from the calibration curve.

Vf = Final prepared (absorbing solution) volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28W7:

If C    = 0.7255 ppb

Vf = 100 mL

W  =  0.51g

S    = 0.792(79.2/100)

DF = 1

Concentration (mg/kg) =     0.7255 x         100          x 1 / 1000

0.51 x 0.792

=  0.1796 mg/kg

=   0.18 mg/kg (Reported Result with Signification)

Calculation  for CN Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg:

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument response in μg/L CN from the calibration curve.

Vf = Final prepared (absorbing solution) volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28W8:

If C    = 33.6234 ppb

Vf = 50 ml

W  = 1.04 g

S    = 0.784(78.4/100)

DF = 1

Concentration (mg/kg) =     33.6234 x          50            x 1 / 1000

1.04 x 0.784

=  2.0618 mg/kg

=  2.1 mg/kg (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements except for Antimony, Arsenic, Cadmium, Selenium, Thallium. Duplicate sample did meet requirements. Serial Dilution did meet requirements.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Antimony | 159Tb |
| Arsenic | 89Y |
| Barium | 159Tb |
| Beryllium | 6Li |
| Cadmium | 159Tb |
| Chromium | 45Sc |
| Cobalt | 45Sc |
| Copper | 45Sc |
| Lead | 209Bi |
| Manganese | 45Sc |
| Nickel | 45Sc |
| Selenium | 89Y |
| Silver | 159Tb |
| Thallium | 209Bi |
| Vanadium | 45Sc |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer