| CHEMTECH PROJECT NUMBER: P5261 | MATRIX: Water |
| --- | --- |
| METHOD: 8270-Modified/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID BN035731.D met the requirements except for 2!4!6-Tribromophenol, The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035747.D met the requirements except for 2,4,6-Tribromophenol,2-Fluorobiphenyl,2-Fluorophenol,Nitrobenzene-d5,Phenol-d6 and Terphenyl-d14 , The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035749.D met the requirements except for Nitrobenzene-d5, The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035845.D met the requirements except for Nitrobenzene-d5, The failure compound not associated with the client parameters list, therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for RE122D1-20241209 [Terphenyl-d14 - 163%], RE122D2-20241209 [Terphenyl-d14 - 196%], RE122D3-20241209 [Terphenyl-d14 - 152%], RE103D1-20241209 [Terphenyl-d14 - 146%], RE103D1-20241209DL [Terphenyl-d14 - 163%], RE103D2-20241209 [Terphenyl-d14 - 146%], RE134D1-20241209 [Terphenyl-d14 - 140%], DUP-01-20241209 [Terphenyl-d14 - 186%], RE134D3-20241210 [Terphenyl-d14 - 136%], RE134D3-20241210DL [Terphenyl-d14 - 148%], RE134D4-20241210 [Terphenyl-d14 - 139%], GM3-RW3-MW1-20241210 [Terphenyl-d14 - 159%], RE125D1-20241210 [Terphenyl-d14 - 142%], RE125D1-20241210MS [Terphenyl-d14 - 146%], RE125D1-20241210MSD [Terphenyl-d14 - 147%], RE125D2-20241210 [Terphenyl-d14 - 148%], RE125D2-20241210DL [Terphenyl-d14 - 155%], RE125D3-20241210 [Terphenyl-d14 - 139%], RE120D1-20241210 [Terphenyl-d14 - 134%], RE120D3-20241210DL [2-Fluorobiphenyl - 113%, Nitrobenzene-d5 - 142%, Terphenyl-d14 - 175%], DUP-02-20241210 [Terphenyl-d14 - 145%], PB165596BS [Nitrobenzene-d5 - 115%], The failure surrogates not associated with the client parameters list, therefore no corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MSD {P5261-14MSD} with File ID: BN035738.D recoveries met the acceptable requirements except for 1,4-Dioxane[50%], marginally low no corrective action is needed. The Blank Spike met requirements for all samples . The RPD for {P5261-14MSD} with File ID: BN035738.D met criteria except for 1,4-Dioxane[40%], due to difference in results of MS and MSD . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The sample # RE125D1-20241210MS and RE125D1-20241210MSD is failing for 1,4-Dioxane, and the original sample(RE125D1-20241210) is reported with M flag for this compounds

Samples RE103D1-20241209, RE134D3-20241210, RE125D2-20241210 and RE120D3-20241210 were diluted due to high concentrations.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date