SDG NARRATIVE

### USEPA

### SDG # ME28Y6

### CASE # 51847

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # P5266

A. Number of Samples and Date of Receipt

18 Soil samples were delivered to the laboratory intact on 12/12/2024

B. Parameters

# Test requested for  Metals CLP12= Aluminum, Calcium, Iron, Magnesium, Potassium, Sodium & Mercury, Cyanide.

# Test requested for Metals CLP MS FULL = Antimony, Arsenic, Barium, Beryllium, Cadmium, Chromium, Cobalt, Copper, Lead, Manganese, Nickel, Selenium, Silver, Thallium, Vanadium, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.3oC, 2.0oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1 : A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: SDGs E28Y6,E28Y9 and ME28Y6 require Laboratory QC but no sample was designated on the COC. The laboratory selected samples E28Y9 and ME28Y9 for Laboratory QC of ARO, SVOA, SVOA SIM, PEST, ICP-MS, ICP-AES, CN and Hg analysis and confirmed these samples are not blank, rinsate or PT samples.

Issue 3: There is no extra volume for soil VOA Laboratory QC and the laboratory would like to proceed without Laboratory QC for soil VOA analysis.

E. Corrective Action taken for above:

Resolution 1 : To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: Per SFAM01.1 Exhibit A, Section 5.5.4.1., the laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 3: Per Region 5, the laboratory will note the issue in the SDG Narrative and proceed without Laboratory QC for soil VOA analysis.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

Inter Element correction factors (IECs) are determined annually and correction factor are applied during

ICP-AES analysis.

G. Calculation:

Calculation for ICP-AES Soil Sample:

Conversion of Results from mg/L or ppm to mg/kg (Dry Weight Basis):

Concentration (mg/kg) =     C  x      Vf      x  DF

W x S

Where,

C   = Instrument value in ppm (The average of all replicate exposures)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28Y6 For Aluminum:

If C    = 70.83167 ppm

Vf = 100 ml

W  = 1.34g

S    = 0.839(83.9/100)

DF = 1

Concentration (mg/kg)   =  70.83167  x        100          x 1

1.34  x 0.839

=    6300.292  mg/kg

=  6300   mg/kg (Reported Result with Signification)

Calculation for ICP-MS Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28Y6 For Arsenic :

If C  =  15.39 ppb

Vf = 500 ml

W  = 1.27 g

S    = 0.839 (83.9/100)

DF = 1

Concentration (mg/kg) =   15.39   x        500       x 1 / 1000

1.27 x 0.839

=  7.22179. mg/kg

=  7.2  mg/kg (Reported Result with Signification)

Calculation for Hg Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument response in μg/L from the calibration curve.

Vf = Final prepared (absorbing solution) volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28Y6:

If C    =  0.1154 ppb

Vf = 100 mL

W  =  0.51 g

S    = 0.839(83.9/100)

DF = 1

Concentration (mg/kg) =   0.1154  x         100          x 1 / 1000

0.51 x 0.839

=  0.02696 mg/kg

=   0.027 mg/kg (Reported Result with Signification)

Calculation  for CN Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg:

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument response in μg/L CN from the calibration curve.

Vf = Final prepared (absorbing solution) volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample ME28Y8:

If C    = 10.3673 ppb

Vf = 50 ml

W  = 1.02 g

S    = 0.827(82.7/100)

DF = 1

Concentration (mg/kg) =   10.3673  x        50         x 1 / 1000

1.02 x 0.827

= 0.61451 mg/kg

=  0.62 mg/kg (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements except for Antimony, Arsenic, Cadmium, Selenium, Thallium, Zinc. Duplicate sample did meet except for Mercury. Serial Dilution did meet requirements.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Antimony | 159Tb |
| Arsenic | 89Y |
| Barium | 159Tb |
| Beryllium | 6Li |
| Cadmium | 159Tb |
| Chromium | 45Sc |
| Cobalt | 45Sc |
| Copper | 45Sc |
| Lead | 209Bi |
| Manganese | 45Sc |
| Nickel | 45Sc |
| Selenium | 89Y |
| Silver | 159Tb |
| Thallium | 209Bi |
| Vanadium | 45Sc |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer