| CHEMTECH PROJECT NUMBER: P5270 | MATRIX: Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF112124.M) for Benzoic acid, Hexachlorocyclopentadiene, 2,4-Dinitrophenol, this compound is passing on Linear Regression. The % RSD is greater than 20% in the Initial Calibration (8270-BF121624.M) for 2,4-Dinitrophenol, 4-Nitrophenol , this compound is passing on Linear Regression and Hexachlorocyclopentadiene is passing on Quadratic regression The Continuous Calibration File ID BF140838.D met the requirements except for 2,4-Dinitrophenol,4,6-Dinitro-2-methylphenol and 4-Nitrophenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken.. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for MW14-20241212DL2 [2,4,6-Tribromophenol - 32%, 2-Fluorobiphenyl - 44%, Nitrobenzene-d5 - 32%, Phenol-d6 - 6% and Terphenyl-d14 - 47%]. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {PB165623BS} with File ID: BF140876.D met requirements for all samples except for 1,2,4,5-Tetrachlorobenzene[102%], 2,2-oxybis(1-Chloropropane)[111%], 4-Bromophenyl-phenylether[104%], Atrazine[122%], bis(2-Chloroethyl)ether[108%], bis(2-Ethylhexyl)phthalate[111%], Butylbenzylphthalate[112%], Indeno(1,2,3-cd)pyrene[106%] and Isophorone[105%], marginally high therefore no further corrective action was taken. The Blank Spike Duplicate for {PB165623BSD} with File ID: BF140877.D met requirements for all samples except for 1,1-Biphenyl[106%], 1,2,4,5-Tetrachlorobenzene[107%], 2,2-oxybis(1-Chloropropane)[105%], 4-Bromophenyl-phenylether[107%], Acenaphthylene[111%], Acetophenone[107%], Anthracene[113%], Atrazine[136%], Benzo(a)anthracene[110%], bis(2-Chloroethyl)ether[104%], bis(2-Ethylhexyl)phthalate[122%], Butylbenzylphthalate[113%], Dibenz(a,h)anthracene[125%], Dimethylphthalate[108%], Di-n-butylphthalate[109%], Indeno(1,2,3-cd)pyrene[128%], Isophorone[108%], N-Nitrosodiphenylamine[115%] and Pyrene[106%], marginally high therefore no further corrective action was taken. . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for MW12-20241212, MW14-20241212 and MW14-20241212, due to high concentration of compounds this sample require dilution therefore sample was reanalyzed with dilution and reported . |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples MW12-20241212, MW12-20241212DL, MW14-20241212 and MW14-20241212DL were diluted due to high concentrations.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date