| CHEMTECH PROJECT NUMBER: P5270 | MATRIX: Water |
| --- | --- |
| METHOD: 8260-Low |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82N112224W.M) for Methyl Acetate compound is passing on Linear Regression. The Continuous Calibration File ID VN085189.D met the requirements except for Carbon Tetrachloride, Tetrachloroethene failing marginally high while, Bromomethane, Dichlorodifluoromethane, Methylcyclohexane, Styrene, and Vinyl Chloride failing high and associated sample required dilution and dilution sample analyzed under passing CCAL therefore no corrective action taken. The Continuous Calibration File ID VN085218.D met the requirements except for 2-Hexanone,4-Methyl-2-Pentanone,Bromoform,Bromomethane,Dibromochloromethane,Dichlorodifluoromethane and Styrene failing high but no positive hit in associated samples therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {VN1213WBS01} with File ID: VN085192.D met requirements for all samples except for Dichlorodifluoromethane[128%],failing high but no positive hit in associated samples therefore no corrective action taken, The Blank Spike Duplicate for {VN1213WBSD01} with File ID: VN085193.D met requirements for all samples except for 1,2-Dibromoethane[112%], Bromoform[112%], Bromomethane[137%], Dibromochloromethane[113%] and Dichlorodifluoromethane[137%],failing high but no positive hit in associated samples therefore no corrective action taken, The Blank Spike for {VN1217WBS02} with File ID: VN085222.D met requirements for all samples except for Bromomethane[126%], Dichlorodifluoromethane[137%],failing high but no positive hit in associated samples therefore no corrective action taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Sample MW12-20241212 was diluted due to high concentration.

sample#MW14-20241212 Initially analyzed but having carry over from above sample as a corrective action sample reanalyzed but required further dilution but now no more vials for reanalysis Therefore sample reported with E flag.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date