| CHEMTECH PROJECT NUMBER: P5279 | MATRIX: Solid |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82Y121724S.M) for Acetone, Chloroform these compounds are passing on Linear Regression. The Continuous Calibration File ID VY020664.D met the requirements except for Bromomethane failing marginally low therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike for {VY1220SBS01} with File ID: VY020666.D met requirements for all samples except for 1,1,2-Trichlorotrifluoroethane[80%],Chloroethane[68%] failing marginally low while 1,2-Dibromo-3-Chloropropane[175%] and t-1,3-Dichloropropene[162%] are failing high but no positive hit in associate sample therefore no corrective action taken. The Blank Spike Duplicate for {VY1220SBSD01} with File ID: VY020667.D met requirements for all samples except for 1,1,2-Trichlorotrifluoroethane[71%], 1,1-Dichloroethane[70%], 1,1-Dichloroethene[70%], Acetone[56%], Bromochloromethane[79%], Methyl Acetate[59%], Methyl tert-butyl Ether[72%] and trans-1,2-Dichloroethene[72%] are failing low therefore as corrective action lab reanalyzed sample but did not purge therefore lab reported sample under failing blank spike duplicate, Also associate CCAL is passing for all analyte. The RPD for {VY1220SBSD01} with File ID: VY020667.D met criteria except for 1,1-Dichloroethane[42%], 1,2-Dibromo-3-Chloropropane[49%], Carbon disulfide[22%], Methyl Acetate[25%], Methyl tert-butyl Ether[33%], Methylene Chloride[28%], t-1,3-Dichloropropene[40%] and trans-1,2-Dichloroethene[40%] due to difference in results of BS and BSD. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

As per special requirement for this project form-1 are reported in mg/kg.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date