| CHEMTECH PROJECT NUMBER: P5291 | MATRIX: Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF121624.M) for 2,4-Dinitrophenol, 4-Nitrophenol, these compounds are passing on Linear Regression and Hexachlorocyclopentadiene is passing on Quadratic regression The Continuous Calibration File ID BF140961.D met the requirements except for 4-Nitrophenol,Benzo(g,h,i)perylene,Dibenzo(a,h)anthracene,Hexachlorocyclopentadiene and Indeno(1,2,3-cd)pyrene, The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P5291-08MS} with File ID: BF140912.D recoveries met the requirements for all compounds except for 1,4-Dioxane[34%], Hexachloroethane[122%] and Pentachlorophenol[140%], due to matrix interference, no corrective action is required. The MSD {P5291-09MSD} with File ID: BF140913.D recoveries met the acceptable requirements except for 1,2,4,5-Tetrachlorobenzene[103%], 1,4-Dioxane[34%], 2,3,4,6-Tetrachlorophenol[124%], Hexachlorobenzene[129%], Hexachloroethane[134%], Nitrobenzene[114%] and Pentachlorophenol[160%], due to matrix interference no corrective action is required. The Blank Spike for {PB165682BS} with File ID: BF140902.D met requirements for all samples except for 1,1-Biphenyl[103%], 2,2-oxybis(1-Chloropropane)[115%], 2-Methylphenol[111%], 3+4-Methylphenols[113%], 4-Bromophenyl-phenylether[108%], Acenaphthylene[118%], Acetophenone[105%], Anthracene[111%], Atrazine[134%], Benzo(a)anthracene[112%], Benzo(k)fluoranthene[124%], bis(2-Chloroethoxy)methane[113%], bis(2-Chloroethyl)ether[107%], bis(2-Ethylhexyl)phthalate[112%], Butylbenzylphthalate[123%], Dibenz(a,h)anthracene[124%], Dimethylphthalate[116%], Di-n-butylphthalate[109%], Indeno(1,2,3-cd)pyrene[127%], Isophorone[112%], N-Nitroso-di-n-propylamine[113%], N-Nitrosodiphenylamine[111%] and Pyrene[115%], The associate samples have no positive hit for these compounds therefore no corrective action was taken. The RPD for {P5291-09MSD} with File ID: BF140913.D met criteria except for Benzaldehyde[21%], due to matrix interference no corrective action is required. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples MW2-20241213, MW3-20241213 were diluted due to high concentrations.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date