| CHEMTECH PROJECT NUMBER: P5291 | MATRIX: Water |
| --- | --- |
| METHOD: 8260-Low |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82N112224W.M) for Methyl Acetate compound this compound is passing on Linear Regression. The %RSD is greater than 20% in the Initial Calibration method (82D121824W.M) for Chloroethane this compound is passing on Linear Regression. The Continuous Calibration File ID VN085218.D met the requirements except for 2-Hexanone,4-Methyl-2-Pentanone, Bromoform, Bromomethane, Dibromochloromethane, Dichlorodifluoromethane and Styrene are failing high and associate sample having hit of 4-Methyl-2-Pentanone but below CRQL therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P5291-08MS} with File ID: VN085240.D recoveries met the requirements for all compounds except for Benzene[180%], Dichlorodifluoromethane[138%], Ethyl Benzene[140%], m/p-Xylenes[160%], o-Xylene[140%] and Toluene[125%] due to matrix interference. The MSD {P5291-09MSD} with File ID: VN085241.D recoveries met the acceptable requirements except for 2-Hexanone[132%], Benzene[160%], Dichlorodifluoromethane[141%], Ethyl Benzene[140%], m/p-Xylenes[150%], Methylcyclohexane[124%], o-Xylene[140%] and Toluene[125%] due to matrix interference. The Blank Spike for {VN1217WBS02} with File ID: VN085222.D met requirements for all samples except for Bromomethane[126%], Dichlorodifluoromethane[137%] are failing high but no positive hit in associate sample therefore no corrective action taken. The Blank Spike for {VN1218WBS01} with File ID: VN085255.D met requirements for all samples except for 2-Hexanone[120%] are failing high but no positive hit in associate sample therefore no corrective action taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples MW2-20241213, MW3-20241213, MW6-20241213 and MW6-20241213-A were diluted due to high concentrations.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date