284 Sheffield Street, Mountainside, NJ 07092 Phone: 908 789 8900 Fax: 908 789 8922

CASE NARRATIVE

### Nobis Group

### Project Name: Raymark Superfund Site

### Project # N/A

### Chemtech Project # P5306

Test Name: SPLP MetalGroup2,SPLP MetalGroup3,SPLP Mercury

A. Number of Samples and Date of Receipt:

18 Solid samples were received on 12/17/2024.

B. Parameters:

According to the Chain of Custody document, the following analyses were requested:

Cyanide, Herbicide Group1, Mercury, Metals Group6, Metals ICP-TAL, METALS-TAL, PCB, Pesticide-TCL, SPLP Extraction, SPLP Mercury, SPLP MetalGroup2, SPLP MetalGroup3, SPLP MetalGroup6, SVOCMS Group3 and VOCMS Group3. This data package contains results for SPLP MetalGroup2,SPLP MetalGroup3,SPLP Mercury.

C. Analytical Techniques:

The analysis of SPLP MetalGroup2,SPLP MetalGroup3 was based on method 6020B, digestion based on method 3010A (Water). The analysis of SPLP Mercury was based on method 7470A and digestion was based on method 7470A (Water).

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Blank Spike met requirements for all samples.

The Duplicate (OU4-VSL-07-121224DUP) analysis met criteria for all samples except for Lead due to sample matrix interference.

The Matrix Spike (OU4-VSL-07-121224MS) analysis met criteria for all samples except for Cadmium, Silver and Thallium due to Chemical Interference during Digestion Process.

The Matrix Spike Duplicate (OU4-VSL-07-121224MSD) analysis met criteria for all samples except for Cadmium and Silver due to Chemical Interference during Digestion Process.

The Blank analysis did not indicate the presence of lab contamination.

The Calibration met the requirements.

The Serial Dilution met the acceptable requirements.

E. Additional Comments:

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag, Cd, Sn, Sb, Ba, Tl, Pb, U are being analyzed with Non-Collision Cell. Helium gas is used for the Collision Cell analysis.

P5306- SPLP all samples are diluted 5X dilution as straight analysis because of high and pure acid concentration of two acids which can cause drastic damage to the instrument.

Internal standard 89Y(1) was out Side qc limit for sample P5306-02 and its Qcs, P5306-08, and P5306-10 in Original so for this sample affected parameters are reported from 10X Dilution.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________