| CHEMTECH PROJECT NUMBER: P5317 | MATRIX: Water |
| --- | --- |
| METHOD: 8270-Modified/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID BN035731.D met the requirements except for 2,4,6-Tribromophenol, The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035747.D met the requirements except for 2,4,6-Tribromophenol,2-Fluorobiphenyl,2-Fluorophenol,Nitrobenzene-d5,Phenol-d6 and Terphenyl-d14, The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035749.D met the requirements except for Nitrobenzene-d5, The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035776.D met the requirements except for 2,4,6-Tribromophenol,The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035794.D met the requirements except for 2,4,6-Tribromophenol, The failure compound not associated with the client parameters list, therefore no corrective action was taken. The Continuous Calibration File ID BN035845.D met the requirements except for Nitrobenzene-d5, The failure compound not associated with the client parameters list, therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for RE114D2-20241216 [Terphenyl-d14 - 155%], RW4-20241216 [Terphenyl-d14 - 143%], TT191D1-20241216 [Terphenyl-d14 - 152%], TT191D2-20241216 [ Terphenyl-d14 - 157%], TT163S1-20241216 [Nitrobenzene-d5 - 112%, Terphenyl-d14 - 179%], DUP-04-20241216 [2-Fluorobiphenyl - 108%, Terphenyl-d14 - 162%], RW8-MW01S-20241216 [Terphenyl-d14 - 182%], RW8-MW01D1-20241216 [Terphenyl-d14 - 191%], TT158S1-20241217 [Nitrobenzene-d5 - 112%, Terphenyl-d14 - 192%], TT158I1-20241217 [Terphenyl-d14 - 154%], DUP-05-20241217 [Terphenyl-d14 - 163%], TT172S1-20241217 [Terphenyl-d14 - 144%], RW9-MW01S-20241217 [2-Fluorobiphenyl - 112%, Terphenyl-d14 - 157%], RW9-MW01D1-20241217 [Terphenyl-d14 - 142%], RW9-MW01D2-20241217 [Terphenyl-d14 - 155%], RW9-MW01D3-20241217 [Terphenyl-d14 - 162%], DUP-06-20241217 [2-Fluorobiphenyl - 111%, Terphenyl-d14 - 169%], RE132D5-20241217 [Terphenyl-d14 - 134%], RE132D5-20241217MS [Terphenyl-d14 - 147%], RE132D5-20241217MSD [Terphenyl-d14 - 168%], RE132D6-20241217 [Terphenyl-d14 - 175%], RE132D6-20241217DL [2-Fluorobiphenyl - 113%, Nitrobenzene-d5 - 130%, Terphenyl-d14 - 222%], RE132D7-20241217 [Terphenyl-d14 - 138%], RW8-MW01D2-20241217 [2-Fluorobiphenyl - 111%, Nitrobenzene-d5 - 122%, Terphenyl-d14 - 190%], RW8-MW01D3-20241217 [2-Fluorobiphenyl - 109%, Nitrobenzene-d5 - 112%, Terphenyl-d14 - 160%], PB165723BL [Nitrobenzene-d5 - 122%, Terphenyl-d14 - 134%], PB165723BS [2-Fluorobiphenyl - 115%, Nitrobenzene-d5 - 141%, Terphenyl-d14 - 142%], PB165724BL [Nitrobenzene-d5 - 114%, Terphenyl-d14 - 134%], PB165724BS [2-Fluorobiphenyl - 115%, Nitrobenzene-d5 - 125%], PB165724BSD [2-Fluorobiphenyl - 116% and Nitrobenzene-d5 - 140%]. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P5317-19MS} with File ID: BN035783.D recoveries met the requirements for all compounds except for 1,4-Dioxane[195%], Due to being a poor compound therefore no corrective action was taken. The Blank Spike met requirements for all samples . The Blank Spike Duplicate met requirements for all samples . The RPD for {P5317-20MSD} with File ID: BN035784.D met criteria except for 1,4-Dioxane[89%], due to difference in results of MS and MSD . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for RE132D6-20241217DL, The failure Internal Standard not associated with the client parameters list, therefore no corrective action was taken. |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples RE132D5-20241217 was diluted due to bad matrix sample run straight where, End Continuous Calibration missing so based on that analysis lab analyzed Sam# RE132D5-20241217 analyzed with direct 2x dilution.

The Sam# RE132D5-20241217, Samples original analyzed with direct 2x dilution is reported as screening data in miscellaneous data.

The sample # RE132D5-20241217MS and RE132D5-20241217MS is failing for cis-1,2-Dichloroethene and the original sample(RE132D5-20241217) is reported with M flag for this compounds.

Sample RE132D6-20241217 was diluted due to high concentration.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date