| CHEMTECH PROJECT NUMBER: P5362 | MATRIX: TCLP |
| --- | --- |
| METHOD: 8270E/3510/1311 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BF122624.M) for 2,4-Dinitrotoluene, this compound is passing on Linear Regression. The Continuous Calibration File ID BF141047.D met the requirements except for 2,4-Dinitrotoluene, The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for WC-SOIL-20241219MS [Terphenyl-d14 - 126%], due to matrix interference no corrective action is required. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P5362-02MS} with File ID: BF141050.D recoveries met the requirements for all compounds except for 2,4,5-Trichlorophenol[114%] and 2,4,6-Trichlorophenol[114%] due to matrix interference no corrective action is required. The Blank Spike for {PB165894BS} with File ID: BF141021.D met requirements for all samples except for 2,4,5-Trichlorophenol[107%], Hexachlorobenzene[109%], The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date