SDG NARRATIVE

### USEPA

### SDG # MJNKA0

### CASE # 51821

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # P5363

A. Number of Samples and Date of Receipt

11 Soil 03 Water samples was delivered to the laboratory intact on 12/20/2024

B. Parameters

# Test requested for Metals CLP4 MS = Arsenic, Copper, Lead, Zinc.

Test requested for Metals CLP MS-CLP4  = Arsenic, Copper, Lead, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.8oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1 : A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

E. Corrective Action taken for above:

Resolution 1 : To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory

will note the issue in the SDG Narrative and proceed with the analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

G. Calculation:

Calculation for ICP-MS Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample MJNKA0 For Arsenic:

If C  =  500.96 ppb

Vf = 500 ml

W  = 1.25 g

S    = 0.798(79.8/100)

DF = 1

Concentration (mg/kg)  =    500.96 x        500          x 1 / 1000

1.25 x 0.798

=  251.1077 mg/kg

=  250 mg/kg (Reported Result with Signification)

Calculation for ICP-MS Water Sample:

Concentration or Result (μg/L) =     C  x    Vf    x  DF

Vi

Where,

C    = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

Vi = Initial aliquot amount (mL) (Sample amount taken in prep)

DF = Dilution Factor

Example Calculation For Sample MJNLE3 For Arsenic:

If C    = 0.18 ppb

Vf = 50 ml

Vi  = 50 ml

DF = 1

Concentration or Result (μg/L) =      0.18 x   50   x 1

50

=   0.18 μg/L

=   0.18  μg/L  (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet

requirements. Duplicate sample did meet requirements except for Arsenic, Lead. Serial Dilution did meet

requirements.

Internal standard 209Bi(1)was out Side qc limit for sample MJNKA5 in Original so for this sample affected parameters are reported from 2X ilution.

Some samples have % solids results less than 50% but more than 30%.  Please see below table for detail. Laboratory has processed these samples according to the SFAM01.1 SOW, Exhibit D, sections 10.1.1.8.

| EPA Sample ID | % Solid |
| --- | --- |
| MJNKJ0 | 49.1 |

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Arsenic | 89Y |
| Copper | 45Sc |
| Lead | 209Bi |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer