| CHEMTECH PROJECT NUMBER: P5371 | MATRIX: Solid |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82N121824W.M) for Chloroethane this compound is passing on Linear Regression. The Continuous Calibration File ID VN085286.D met the requirements except for 2-Butanone failing low therefore all associated samples were reanalyzed to confirm the failure and both run reported. The Continuous Calibration File ID VN085309.D met the requirements except for 2-Butanone and Acetone are failing low but all associated samples are analyzed for confirmation of CCAL failure of first analysis therefore both run reported. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for SPLP-C10-2-M [Dibromofluoromethane - 66%], SPLP-C10-3-M [Dibromofluoromethane - 65%], SPLP-C11-1-M [Dibromofluoromethane - 61%], SPLP-C11-2-M [Dibromofluoromethane - 63%], SPLP-C11-3-M [Dibromofluoromethane - 64%], SPLP-C11-3-MRE [Dibromofluoromethane - 65%], SPLP-C12-1-M [Dibromofluoromethane - 63%], SPLP-C12-1-MRE [Dibromofluoromethane - 60%], SPLP-C12-2-M [Dibromofluoromethane - 62%], SPLP-C12-2-MRE [Dibromofluoromethane - 61%], SPLP-C12-3-M [Dibromofluoromethane - 64%] and SPLP-C12-3-MRE [Dibromofluoromethane - 65%], these compounds did not meet the NJDKQP criteria but met the in-house criteria . |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Trip Blank was not provided with this set of samples.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date