| CHEMTECH PROJECT NUMBER: P5374 | MATRIX: Water |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82x121124w.M) for t-1,3-Dichloropropene, Dibromochloromethane, Bromoform These compounds are passing on Quadratic Regression. The Continuous Calibration File ID VX044466.D met the requirements except for Methylcyclohexane and Tetrachloroethene failing high but no positive hit in associated samples therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for BP-VPB-190A-GW-748-750 [Toluene-d8 - 120%], BP-VPB-190A-GW-748-750MS [1,2-Dichloroethane-d4 - 79%], BP-VPB-190A-GW-748-750MSD [1 and2-Dichloroethane-d4 - 80%] Surrogate failure for MS-MSD confirms with original sample. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {P5374-05MS} with File ID: VX044482.D recoveries met the requirements for all compounds except for Trichloroethene[140%] due to matrix interference. The MSD {P5374-06MSD} with File ID: VX044483.D recoveries met the acceptable requirements except for Trichloroethene[160%]due to matrix interference. The sample # BP-VPB-190A-GW-748-750MS and BP-VPB-190A-GW-748-750MSD is failing for Trichloroethene and the original sample(BP-VPB-190A-GW-748-750) is reported with M flag for this compound. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Sample #BP-VPB-190A-TB-20241218 , BP-VPB-190A-EB-20241218 have the concentration of target compound below Method detection limits, therefore it is not reported as Hit in Form1

The laboratory certifies that the all-electronic diskette deliverable exactly match the data summary forms (i.e. Form Is).”

The not QT review data is reported in the Miscellaneous.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date