| CHEMTECH PROJECT NUMBER: Q1069 | MATRIX: Solid |
| --- | --- |
| METHOD: 8270E/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for TR-06-1-10-2025MS [2,4,6-Tribromophenol - 22%, Terphenyl-d14 - 45%], TR-06-1-10-2025MSD [2,4,6-Tribromophenol - 25%, Terphenyl-d14 - 52%],due to matrix interference and no corrective action is needed, and RW7B-CARBON-20250109 [2,4,6-Tribromophenol - 6%, 2-Fluorobiphenyl - 9%, 2-Fluorophenol - 5%, Nitrobenzene-d5 - 9%, Phenol-d6 - 5%, Terphenyl-d14 - 8%], RW7B-CARBON-20250109RX [2,4,6-Tribromophenol - 34%, 2-Fluorobiphenyl - 39%, 2-Fluorophenol - 33%, Phenol-d6 - 32% and Terphenyl-d14 - 29%]Failure sample for surrogate was reanalyzed to confirm the failure and both run were reported in Hard Copy. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1068-01MS} with File ID: BF141175.D recoveries met the requirements for all compounds except for 1,4-Dioxane[48%], 2,3,4,6-Tetrachlorophenol[23%], 2,4,6-Trichlorophenol[35%], 2,4-Dinitrophenol[0%], 4,6-Dinitro-2-methylphenol[12%], 4-Nitrophenol[21%] and Pentachlorophenol[24%], due to matrix interference and no corrective action is needed. . The MSD {Q1068-01MSD} with File ID: BF141176.D recoveries met the acceptable requirements except for 1,4-Dioxane[49%], 2,3,4,6-Tetrachlorophenol[26%], 2,4,6-Trichlorophenol[35%], 2,4-Dinitrophenol[0%], 4,6-Dinitro-2-methylphenol[9%], 4-Nitrophenol[25%] and Pentachlorophenol[23%], due to matrix interference and no corrective action is needed. The Blank Spike for {PB166035BS} with File ID: BF141165.D met requirements for all samples except for Hexachlorocyclopentadiene[179%], The associate samples have no positive hit for these compounds therefore no corrective action was taken. The Blank Spike for {PB166061BS} with File ID: BF141172.D met requirements for all samples except for Atrazine[129%], Hexachlorocyclopentadiene[191%],The associate samples have no positive hit for these compounds therefore no corrective action was taken. . |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The laboratory certifies that the all-electronic diskette deliverable exactly match the data

summary forms (i.e. Form Is).”

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date