SDG NARRATIVE

### USEPA

### SDG # MC0B99

### CASE # 51943

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # Q1085

A. Number of Samples and Date of Receipt

05 Water samples were delivered to the laboratory intact on 01/15/2025.

B. Parameters

# Test requested for Metals CLP4 = Calcium, Magnesium, Hardness Total & Mercury.

# Test requested for Metals CLP MS = Aluminum, Antimony, Arsenic, Barium, Beryllium, Cadmium, Calcium, Chromium, Cobalt, Copper, Iron, Lead, Magnesium, Manganese, Nickel, Potassium, Selenium, Silver, Sodium, Thallium, Vanadium, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  3.1oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1: A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: Hardness is not scheduled for this Case, but the COC analysis key indicates that hardness is required.

Issue 3: A 7-day TAT is scheduled for this Case, but the COC analysis key indicates a 14-day TAT.

E. Corrective Action taken for above:

Resolution 1: To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: As per Region 3, there are nine aqueous samples scheduled for ICP-AES 1-4 Metals; the nine samples are to be analyzed for Ca, Mg, and hardness. Hardness will be scheduled for the nine

samples. Please note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: As per Region 3, please make note of the issue in the SDG Narrative and proceed with the

analysis of the samples as scheduled (7-Day TAT).

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

Inter Element correction factors (IECs) are determined annually and correction factor are applied during

ICP-AES analysis.

G. Calculation:

Calculation for ICP-AES Water Sample:

Concentration or Result (μg/L) =     C  x    Vf    x  DF x 1000

Vi

Where,

C  = Instrument value in ppm (The average of all replicate exposures)

Vf = Final digestion volume (mL)

Vi = Initial aliquot amount (mL) (Sample amount taken in prep)

DF = Dilution Factor

Example Calculation For Sample MC0B99 For Calcium:

If C    = 68.10030  ppm

Vf = 50 ml

Vi  = 50 ml

DF = 1

Concentration or Result (μg/L) =     68.10030  x   50   x 1 x 1000

50

= 68100.30 μg/L

= 68000 μg/L  (Reported Result with Signification)

Calculation for ICP-MS Water Sample:

Concentration or Result (μg/L) =     C  x    Vf    x  DF

Vi

Where,

C    = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

Vi = Initial aliquot amount (mL) (Sample amount taken in prep)

DF = Dilution Factor

Example Calculation For Sample MC0B99 For Arsenic:

If C    = 0.35 ppb

Vf = 50 ml

Vi  = 50 ml

DF = 1

Concentration or Result (μg/L) =     0.35  x   50   x 1

50

=   0.35 μg/L

=   0.35 μg/L  (Reported Result with Signification)

Calculation for Hg Water Sample:

Concentration or Result (μg/L) =     C  x  DF

Where,

C    = Instrument response in μg/L from the calibration curve.

DF = Dilution Factor

Example Calculation For Sample MC0B99:

If C    = 0.0534 ppb

DF = 1

Concentration or Result (μg/L) =     0.0534 x  1

=    0.0534 μg/L

=     0.053 μg/L  (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements. Duplicate sample did meet requirements. Serial Dilution did meet requirements.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Aluminum | 45Sc |
| Antimony | 159Tb |
| Arsenic | 89Y |
| Barium | 159Tb |
| Beryllium | 6Li |
| Cadmium | 159Tb |
| Calcium | 45Sc |
| Chromium | 45Sc |
| Cobalt | 45Sc |
| Copper | 45Sc |
| Iron | 45Sc |
| Lead | 209Bi |
| Magnesium | 45Sc |
| Manganese | 45Sc |
| Nickel | 45Sc |
| Potassium | 45Sc |
| Selenium | 89Y |
| Silver | 159Tb |
| Sodium | 45Sc |
| Thallium | 209Bi |
| Vanadium | 45Sc |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer