SDG NARRATIVE

### USEPA

### SDG # MH2H28

### CASE # 51822

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # Q1102

A. Number of Samples and Date of Receipt

04 Soil samples were delivered to the laboratory intact on 01/16/2025

B. Parameters

# Test requested for Metals CLP MS FULL = Antimony, Arsenic, Barium, Beryllium, Cadmium, Chromium, Cobalt, Copper, Lead, Manganese , Nickel, Selenium, Silver, Thallium, Vanadium, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.1oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1: A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: The laboratory received soil samples for ICP-MS analysis with a 14 day TAT. However, the COC does not list MA 3105.0. The laboratory would like to confirm if the received samples should be analyzed under MA 3105.0.

E. Corrective Action taken for above:

Resolution 1: To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: Per Region 8, the received samples should be analyzed without MA 3105.0. The laboratory should note the issue in the SDG Narrative and proceed with the analysis of the samples.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

Inter Element correction factors (IECs) are determined annually and correction factor are applied during

ICP-AES analysis.

G. Calculation:

Calculation for ICP-MS Soil Sample:

Conversion of Results from μg /L or ppb to mg/kg :

Concentration (mg/kg) =     C  x      Vf      x  DF / 1000

W x S

Where,

C   = Instrument value in ppb (The average of all replicate integrations)

Vf = Final digestion volume (mL)

W  = Initial aliquot amount (g) (Fraction of Sample amount taken in prep)

S    = % Solids / 100 (Fraction of Percent Solids)

DF = Dilution Factor

Example Calculation For Sample  MH2H28 For Antimony:

If C  =  11.15 ppb

Vf = 500 ml

W  = 1.14 g

S    = 1(100/100)

DF = 1

Concentration (mg/kg) =   11.15  x        500          x 1 / 1000

1.14 x 1

=  4.8904 mg/kg

=  4.9 mg/kg (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements except for Arsenic and Copper. Duplicate sample did meet requirements. Serial Dilution did meet requirements.

Collision cell is being used to remove potential interferences. The analytes Na, Mg, Al, K, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, As are being analyzed with collision cell and analytes Be, B, Ca, Ti, Se, Sr, Zr, Mo, Ag,

Cd, Sn, Sb, Ba, Tl, Pb, U  are being analyzed with Non-Collision Cell. Helium gas is used for the

Collision Cell analysis.

Internal Standard Association for ICP-MS analysis.

| Target Analyte | Associated Internal Standard |
| --- | --- |
| Antimony | 159Tb |
| Arsenic | 89Y |
| Barium | 159Tb |
| Beryllium | 6Li |
| Cadmium | 159Tb |
| Chromium | 45Sc |
| Cobalt | 45Sc |
| Copper | 45Sc |
| Lead | 209Bi |
| Manganese | 45Sc |
| Nickel | 45Sc |
| Selenium | 89Y |
| Silver | 159Tb |
| Thallium | 209Bi |
| Vanadium | 45Sc |
| Zinc | 45Sc |

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer