| CHEMTECH PROJECT NUMBER: Q1194 | MATRIX: Solid, Water |
| --- | --- |
| METHOD: 6010D,7470A,7471B |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Calibration Summary met criteria. |  |  |  |
| 2. | ICP Interference Check Sample Results Summary Submitted. |  |  |  |
| 3. | Serial Dilution Summary (if applicable) Submitted. The Serial Dilution (B-113-SB02L) met criteria for all samples except for Calcium, Chromium, Iron, Magnesium and Manganese due to sample matrix interference. |  |  |  |
| 4. | Laboratory Control Sample Summary (if applicable) Submitted. |  |  |  |
| 5. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Met Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Matrix Spike (B-113-SB02MS) analysis met criteria for all samples except for Beryllium, Potassium and Selenium due to Chemical Interference during Digestion process. The Matrix Spike (EBMS) analysis met criteria for all samples except for Manganese due to Chemical Interference during Digestion Process. The Matrix Spike Duplicate (B-113-SB02MSD) analysis met criteria for all samples except for Barium, Beryllium, Chromium, Copper, Potassium, Selenium and Vanadium due to Chemical Interference during Digestion Process. The Matrix Spike Duplicate (EBMSD) analysis met criteria for all samples except for Manganese due to Chemical Interference during Digestion Process. |  |  |  |
| 7. | Sample Duplicate Analysis Met QC Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Duplicate (B-113-SB02DUP) analysis met criteria for all samples except for Cadmium due to sample matrix interference. |  |  |  |
| 8. | Digestion Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

_______________________________________				__________________________

QA REVIEW								Date