| CHEMTECH PROJECT NUMBER: Q1194 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270E/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The % RSD is greater than 20% in the Initial Calibration (8270-BG020325.M) for Nitrobenzene, 2,4-Dichlorophenol,Caprolactam, 2,4,6-Tribromophenol, 2,4,6-Trichlorophenol, 2,4,5-Trichlorophenol , 3-Nitroaniline,, 4-Nitrophenol, 2,3,4,6-Tetrachlorophenol,4-Nitroaniline, Pentachlorophenol, Butylbenzylphthalate , Bis(2-ethylhexyl)phthalate, Di-n-octyl phthalate, this compound is passing on Linear Regression and 2-Nitrophenol, 2-Nitroaniline , 2,6-Dinitrotoluene, 2,4-Dinitrophenol, 2,4-Dinitrotoluene, 4,6-Dinitro-2-methylphenol, is passing on Quadratic regression The Continuous Calibration File ID BF141336.D met the requirements except for 4-Nitrophenol and Pentachlorophenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken. The Continuous Calibration File ID File ID BF141410.D met the requirements except for 4,6-Dinitro-2-methylphenol, The associate samples have no positive hit for these compounds therefore no corrective action was taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1194-04MS} with File ID: BF141352.D recoveries met the requirements for all compounds except for 3,3-Dichlorobenzidine[58%], 3-Nitroaniline[54%], 4-Chloroaniline[17%], Benzaldehyde[13%], Benzo(g,h,i)perylene[58%], Dibenz(a,h)anthracene[69%], Indeno(1,2,3-cd)pyrene[69%] and Pyrene[69%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. The MSD {Q1194-04MSD} with File ID: BF141353.D recoveries met the acceptable requirements except for 3,3-Dichlorobenzidine[62%], 3-Nitroaniline[54%], 4-Chloroaniline[23%], Benzaldehyde[13%], Benzo(g,h,i)perylene[58%], Dibenz(a,h)anthracene[69%], Indeno(1,2,3-cd)pyrene[69%] and Pyrene[65%], these compounds did not meet the NJDKQP criteria but met the in-house criteria. The Blank Spike for {PB166294BS} with File ID: BF141439.D met requirements for all samples except for 3,3-Dichlorobenzidine[47%], 3-Nitroaniline[53%], 4-Chloroaniline[42%], Benzaldehyde[12%], these compounds did not meet the NJDKQP criteria but met the in-house criteria, while Hexachlorocyclopentadiene[185%], this compound did not meet the NJDKQP criteria and in-house criteria, the associate samples have no positive hit for these compounds therefore no corrective action was taken. The Blank Spike for {PB166339BS} with File ID: BF141442.D met requirements for all samples except for 3,3-Dichlorobenzidine[46%], 4-Chloroaniline[39%],3-Nitroaniline[53%], Benzaldehyde[13%], these compounds did not meet the NJDKQP criteria but met the in-house criteria, while Hexachlorocyclopentadiene[190%], this compound did not meet the NJDKQP criteria and in-house criteria, the associate samples have no positive hit for these compounds therefore no corrective action was taken. The Blank Spike Duplicate for {PB166339BSD} with File ID: BF141443.D met requirements for all samples except for 3,3-Dichlorobenzidine[47%], 3-Nitroaniline[54%], 4-Chloroaniline[37%], Benzaldehyde[12%], these compounds did not meet the NJDKQP criteria but met the in-house criteria, Hexachlorocyclopentadiene[190%], this compound did not meet the NJDKQP criteria and in-house criteria, the associate samples have no positive hit for these compounds therefore no corrective action was taken. The RPD for {Q1194-04MSD} with File ID: BF141353.D met criteria except for 4-Chloroaniline[30%], this compound did not meet the NJDKQP criteria but met the in-house criteria. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date