| CHEMTECH PROJECT NUMBER: Q1199 | MATRIX: Water |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration File ID VN085526.D met the requirements except for 2-Hexanone and Styrene failing high but associated sample having hit below CRQL therefore no corrective action taken. |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike met requirements for all samples . The Blank Spike Duplicate met requirements for all samples |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

This data package has been revised due to client ID changed for sample#07 as per client request.

The Sample #BP-VPB-192-GW-220-222 have the concentration of  target compound  below Method detection limits, therefore it is not reported as Hit in Form1.

The laboratory certifies that the all-electronic diskette deliverable exactly match the data summary forms (i.e. Form Is).”

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

The not QT review data is reported in the Miscellaneous.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date