SDG NARRATIVE

### USEPA

### SDG # MBHHA8

### CASE # 51979

CONTRACT # 68HERH20D0011

SOW# SFAM01.1

### LAB NAME: Alliance Technical Group, LLC

### LAB CODE: ACE

### LAB ORDER ID # Q1315

A. Number of Samples and Date of Receipt

01 Water samples was delivered to the laboratory intact on 02/06/2025.

B. Parameters

# Test requested for Metals CLP FULL = Aluminum, Antimony, Arsenic, Barium, Beryllium, Cadmium, Calcium, Chromium, Cobalt, Copper, Iron, Lead, Magnesium, Manganese, Nickel, Potassium, Selenium, Silver, Sodium, Thallium, Vanadium, Zinc.

# C. Cooler Temp

Indicator Bottle: Presence/Absence

Cooler:  2.6oC

# D. Detail Documentation (related to Sample Handling

Shipping, Analytical Problem, Temp of Cooler etc):

Issue 1 : A "P" or "M" prefix was listed at the beginning of a CLP sample ID.

Issue 2: Laboratory QC is scheduled for soil and water ICP-AES Metals analysis, but no samples were

Designated for QC. The laboratory would like to use samples MBHHA7, MBHHA8, and MBHHA9 for

laboratory QC. The laboratory confirms that these samples are not blanks, rinsates, or PT samples.

E. Corrective Action taken for above:

Resolution 1 : To maintain COC integrity, ASB requests no changes to the Sample IDs. The laboratory will note the issue in the SDG Narrative and proceed with the analysis of the samples.

Resolution 2: Per SOW SFAM01.1 Exhibit A, Section 5.5.4.1., please note the issue in the SDG Narrative

And proceed with the analysis of the samples using the selected samples for QC.

F. Analytical Techniques:

All analyses were based on CLP Methodology by method SFAM01.1.

Inter Element correction factors (IECs) are determined annually and correction factor are applied during

ICP-AES analysis.

G. Calculation:

Calculation for ICP-AES Water Sample:

Concentration or Result (μg/L) =     C  x    Vf    x  DF x 1000

Vi

Where,

C  = Instrument value in ppm (The average of all replicate exposures)

Vf = Final digestion volume (mL)

Vi = Initial aliquot amount (mL) (Sample amount taken in prep)

DF = Dilution Factor

Example Calculation For Sample MBHHA8 For Aluminum:

If C    = 0.0273325 ppm

Vf = 50 ml

Vi  = 50 ml

DF = 1

Concentration or Result (μg/L) =     0.0273325 x   50   x 1 x 1000

50

=   27.3325   μg/L

=   27 μg/L  (Reported Result with Signification)

H. QA/ QC

Calibrations met requirements.  Interference check met requirements.  Blank analyses did not indicate any

presence of contamination. Laboratory Control sample was within control limits. Spike sample did meet requirements except for Silver. Duplicate sample did meet requirements. Serial Dilution did meet requirements.

I certify that the data package is in compliance with the terms and conditions of the contract both technically and for completeness, for other than the conditions detailed above.  Release of the data contained in this hard copy data package has been authorized by the Laboratory Director or his designee, as verified by the following signature.

Signature__________________________	Name: Nimisha Pandya

Date _____________________________ 	Title: Document Control Officer