CASE NARRATIVE

### USEPA CLP Organics

### Project Name: 51900

### Project # N/A

### Chemtech Project # Q1335

Test Name: VOC-TRACE-SFAM

A. Number of Samples and Date of Receipt:

10 Water samples were received on 02/07/2025.

B. Parameters

According to the Chain of Custody document, the following analyses were requested:

VOC-TRACE-SFAM. This data package contains results for VOC-TRACE-SFAM.

C. Analytical Techniques:

The analysis performed on instrument MSVOA_V were done using GC column DB-624UI 20m 0.18mm 1.0 um. Cat#121-1324UIThe analysis of VOC-TRACE-SFAM was based on method SFAM_Trace.

D. QA/ QC Samples:

The Holding Times were met for all analysis.

The Surrogate recoveries met the acceptable criteria except for

E29C3 [Toluene-d8 - 66%],

E29C4 [2-Hexanone-d5 - 134%].

The Internal Standards Areas met the acceptable requirements.

The Retention Times were acceptable for all samples.

The Blank analysis did not indicate the presence of lab contamination.

The Initial Calibration met the requirements .

The Continuous Calibration File ID VV038662.D met the requirements except for 2-Butanone,Acetone,Bromoform,1,1-Dichloroethene-d2,2-Butanone-d5,Chloroethane-d5 and Vinyl chloride-d3 .

The Continuous Calibration File ID VV038674.D met the requirements except for Bromomethane .

The Continuous Calibration File ID VV038688.D met the requirements except for 2-Hexanone,4-Methyl-2-pentanone and Vinyl chloride-d3 .

The Continuous Calibration File ID VV038695.D met the requirements except for 1!1-Dichloroethene-d2,trans-1!3-Dichloropropene-d4 and Vinyl chloride-d3 .

The Tuning criteria met requirements.

E. Additional Comments:

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

F. Manual Integration Comments:

Please refer to the Manual integration Report included with the Run Logs for information on the manual integrations performed.

I certify that the data package is in compliance with the terms and conditions of the contract, both technically and for completeness, for other than the conditions detailed above. The laboratory manager or his designee, as verified by the following signature has authorized release of the data contained in this hard copy data package.

Signature__________________________