| CHEMTECH PROJECT NUMBER: Q1380 | MATRIX: /Water |
| --- | --- |
| METHOD: 8270-Modified/3510 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications. DFTPP Meet Criteria. (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The Initial Calibration met the requirements . The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for PB166758BS [2-Fluorobiphenyl - 118%], BP-VPB-192-EB-20250212 [Terphenyl-d14 - 146%], BP-VPB-192-GW-780-782 [2-Fluorobiphenyl - 118%], BP-VPB-192-GW-780-782MS [2-Fluorobiphenyl - 110%, Terphenyl-d14 - 135%], BP-VPB-192-GW-780-782MSD [2-Fluorobiphenyl - 113%, Terphenyl-d14 - 138%],failure surrogates are not associated with the client list, as per criteria affected surrogates were passing, therefore no corrective action was taken while, VPB192-HYD-20250214 [2-Methylnaphthalene-d10 - 6%]surrogate failing biased low also Due to the limited volume of this sample, it will not be re-extracted therefore no corrective action taken. |  |  |  |

| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1380-10MS} with File ID: BN036477.D recoveries met the requirements for all compounds except for 1,4-Dioxane[64%] due to matrix interference. The MSD recoveries met the acceptable requirements . The sample # BP-VPB-192-GW-780-782MS is failing for 1,4-Dioxane and the original sample(BP-VPB-192-GW-780-782) is reported with M flag for this compound.. The Blank Spike met requirements for all samples . |  |  |  |
| --- | --- | --- | --- | --- |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 11. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

The laboratory certifies that the all-electronic diskette deliverable exactly match the data summary forms (i.e. Form Is).

Less volume was taken for samples #BP-VPB-192-GW-780-782, BP-VPB-192-GW-780-782MS,BP-VPB-192-GW-780-782MSD at the time of extraction due to Limited volume received.

The Form 6 is not included in the data package because the Initial Calibration was performed using 7 points.

The not QT review data is reported in the Miscellaneous.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date