| CHEMTECH PROJECT NUMBER: Q1380 | MATRIX: /Water |
| --- | --- |
| METHOD: 8260D |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82N021825W.M) for Styrene this compound is passing on Quadratic Regression. The Continuous Calibration met the requirements . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for BP-VPB-192-GW-725-727MS [4-Bromofluorobenzene - 63%, Toluene-d8 - 73%] ,Surrogate fail in Only MS but Parent Sample and MSD are Passing For Surrogate Recoveries therefore no Corrective action was taken. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1380-07MS} with File ID: VY021342.D recoveries met the requirements for all compounds except for 1,2-Dichlorobenzene[72%], 1,3-Dichlorobenzene[67%], 1,4-Dichlorobenzene[69%], Chlorobenzene[76%], Ethyl Benzene[75%], m/p-Xylenes[74%], Methylcyclohexane[46%], o-Xylene[75%], Styrene[51%], Tetrachloroethene[72%] and Trichloroethene[76%], due to Matrix interference. The MSD {Q1380-08MSD} with File ID: VY021343.D recoveries met the acceptable requirements except for 1,1,2,2-Tetrachloroethane[164%], Methyl tert-butyl Ether[127%] and Styrene[69%], due to Matrix interference. Fail in MS-MSD-The sample # BP-VPB-192-GW-725-727MS and BP-VPB-192-GW-725-727MSD are failing for Methyl tert-butyl Ether,Methylcyclohexane,Trichloroethene,Tetrachloroethene,Chlorobenzene,Ethyl Benzene,m/p-Xylenes,o-Xylene,Styrene,1,1,2,2-Tetrachloroethane,1,3-Dichlorobenzene,1,4-Dichlorobenzene and 1,2-Dichlorobenzene the original sample(BP-VPB-192-GW-725-727) is reported with M flag for this compounds. The RPD for {Q1380-08MSD} with File ID: VY021343.D met criteria except for 1,1,1-Trichloroethane[34%], 1,1,2,2-Tetrachloroethane[38%], 1,1,2-Trichloroethane[40%], 1,1,2-Trichlorotrifluoroethane[37%], 1,1-Dichloroethane[32%], 1,1-Dichloroethene[35%], 1,2-Dichlorobenzene[40%], 1,2-Dichloroethane[38%], 1,2-Dichloropropane[35%], 1,3-Dichlorobenzene[38%], 1,4-Dichlorobenzene[40%], 2-Butanone[41%], 2-Hexanone[45%], 4-Methyl-2-Pentanone[45%], Acetone[39%], Benzene[35%], Bromodichloromethane[34%], Bromoform[38%], Bromomethane[33%], Carbon disulfide[32%], Carbon Tetrachloride[35%], Chlorobenzene[34%], Chloroethane[31%], Chloroform[33%], Chloromethane[36%], cis-1,2-Dichloroethene[33%], cis-1,3-Dichloropropene[35%], Dibromochloromethane[38%], Ethyl Benzene[35%], Isopropylbenzene[37%], m/p-Xylenes[37%], Methyl tert-butyl Ether[39%], Methylcyclohexane[43%], Methylene Chloride[33%], o-Xylene[35%], Styrene[30%], t-1,3-Dichloropropene[36%], Tetrachloroethene[34%], Toluene[35%], trans-1,2-Dichloroethene[34%], Trichloroethene[35%], Trichlorofluoromethane[35%] and Vinyl chloride[37%] , due to difference in results of MS and MSD. The Blank Spike for {VY0219SBS01} with File ID: VY021238.D met requirements for all samples except for Chloroethane[146%], Chloromethane[150%] and Vinyl chloride[145%], are failing high but no positive hit in associate samples therefore no corrective action taken. The Blank Spike Duplicate met requirements for all samples. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: The Internal Standards Areas met the acceptable requirements except for BP-VPB-192-GW-725-727MS, BP-VPB-192-GW-725-727MSD.Internal standard fail in only MSD but MS and Parent Sample are Passing For Internal standard Recoveries therefore no Corrective action was taken. |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: The Holding Times were met for all analysis except for BP-VPB-192-GW-725-727MS and BP-VPB-192-GW-725-727MSD,for MS-MSD VIAL A Initially analyzed in sequence VY022025 with in Holding time but internal standard and surrogate failed as well as End CCAL was failing therefore as a corrective action lab analyzed MS-MSD again with VIAL B where Internal standards and surrogate failure confirmed but this analysis is out of Holding time, Therefore VIAL B reported as final and VIAL A reported as screening data in miscellaneous section. |  |  |  |

ADDITIONAL COMMENTS:

The laboratory certifies that the all-electronic diskette deliverable exactly match the data summary forms (i.e. Form Is).”

The not QT review data is reported in the Miscellaneous.

The soil samples  results  are based on a dry weight basis.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <20% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 20% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date