| CHEMTECH PROJECT NUMBER: Q1405 | MATRIX: Water |
| --- | --- |
| METHOD: 8260-Low |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. (Field samples and Method Blanks) |  |  | ü |
| 2. | GC/MS Tuning Specifications BFB Meet Criteria (NOTE THAT THERE ARE DIFFERENT CRITERIA FOR NY ASP CLP, CLP AND NJ) |  |  | ü |
| 3. | GC/MS Tuning Frequency - Performed every 24 hours for 600 series and 12 hours for 8000 Series. |  |  | ü |
| 4. | GC/MS Calibration - Initial Calibration performed before sample analysis and continuing calibration performed within 24 hours of sample analysis for 600 series and 12 hours for 8000 series. |  |  | ü |
| 5. | GC/MS Calibration Requirements. |  |  |  |
|  | The %RSD is greater than 20% in the Initial Calibration method (82N021825W.M) for Ethyl methacrylate, 2-Chloroethyl Vinyl ether, Styrene, p-Isopropyl toluene these compound are passing on Quadratic Regression and Isopropyl Acetate is passing on Linear Regression. The Continuous Calibration File ID VN085802.D met the requirements except for 4-Bromofluorobenzene are failing high but no positive hit in associate samples therefore no corrective action taken.. The Continuous Calibration File ID VN085829.D met the requirements except for 2-Hexanone,4-Methyl-2-Pentanone,Styrene and 4-are failing high but no positive hit in associate sample therefore no corrective action taken. . |  |  |  |
| 6. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 7. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. |  |  |  |
| 8. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The Blank Spike Duplicate for {VN0221WBSD01} with File ID: VN085806.D met requirements for all samples except for 1,2-Dibromoethane[113%], 2-Hexanone[130%], 4-Methyl-2-Pentanone[120%], Bromoform[112%] and Dibromochloromethane[112%] are failing high but no positive hit in associate samples therefore no corrective action taken. The Blank Spike for {VN0224WBS01} with File ID: VN085832.D met requirements for all samples except for 2-Hexanone[120%], 4-Methyl-2-Pentanone[120%], Bromochloromethane[125%] and Bromoform[111%] are failing high but no positive hit in associate sample therefore no corrective action taken. |  |  |  |
| 9. | Internal Standard Area/Retention Time Shift Meet Criteria Comments: |  |  |  |
| 10. | Analysis Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |

ADDITIONAL COMMENTS:

Sample MW-04-021325 was diluted due to high concentration.

Samples for MS/MSD for VOC analysis were not provided with this set of samples. The Blank Spike Duplicate is reported with the data.

Please use %D calculated based on Avg RF and CCRF for all compounds using Average Response Factor when the %RSD value for a compound is <15% for the Initial Calibration curve and use %D calculated based on Amount added and Calculated amount for all compounds using Linear Regression when the %RSD value for a compound is > 15% for the Initial Calibration curve for SW-846 analysis.

_______________________________________				__________________________

QA REVIEW								Date