| CHEMTECH PROJECT NUMBER: Q1433 | MATRIX: Solid |
| --- | --- |
| METHOD: 8082A/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for K084-1A [Decachlorobiphenyl(1) - 746%, Decachlorobiphenyl(2) - 793%], K084-1ADL [Decachlorobiphenyl(1) - 822%, Decachlorobiphenyl(2) - 981%], K084-2A [Decachlorobiphenyl(1) - 310%, Decachlorobiphenyl(2) - 339%], K084-2ADL [Decachlorobiphenyl(1) - 357%, Decachlorobiphenyl(2) - 396%], K084-2ADL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], K084-2B [Decachlorobiphenyl(1) - 909%, Decachlorobiphenyl(2) - 959%], K084-2BDL [Decachlorobiphenyl(1) - 993%, Decachlorobiphenyl(2) - 1074%], K084-2BDL2 [Decachlorobiphenyl(1) - 1152%, Decachlorobiphenyl(2) - 1284%], K084-3A [Decachlorobiphenyl(1) - 1809%, Decachlorobiphenyl(2) - 1855%], K084-3ADL [Decachlorobiphenyl(1) - 2105%, Decachlorobiphenyl(2) - 2223%, Tetrachloro-m-xylene(1) - 148%, Tetrachloro-m-xylene(2) - 157%], K084-3ADL2 [Decachlorobiphenyl(1) - 2369%, Decachlorobiphenyl(2) - 2621%, Tetrachloro-m-xylene(1) - 166%, Tetrachloro-m-xylene(2) - 166%], K084-3B [Decachlorobiphenyl(1) - 177%, Decachlorobiphenyl(2) - 198%], K084-3BDL [Decachlorobiphenyl(1) - 261%, Decachlorobiphenyl(2) - 274%], K084-4A [Decachlorobiphenyl(1) - 231%, Decachlorobiphenyl(2) - 265%], K084-4ADL [Decachlorobiphenyl(1) - 291%, Decachlorobiphenyl(2) - 335%], K084-4ADL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], K084-5AMS [Decachlorobiphenyl(1) - 176%, Decachlorobiphenyl(2) - 205%], K084-5AMSD [Decachlorobiphenyl(1) - 180% and Decachlorobiphenyl(2) - 209%]. Due to high concentration of compounds, this sample required dilution. Therefore, sample was reanalyzed with dilution and reported. |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1434-03MS} with File ID: PO109562.D recoveries met the requirements for all compounds except for AR1016[182%], AR1260[554%] due to matrix interference. The MS {Q1434-02MS} with File ID: PO109688.D recoveries met the requirements for all compounds except for AR1260[309%] due to matrix interference. The MSD {Q1434-03MSD} with File ID: PO109563.D recoveries met the acceptable requirements except for AR1016[182%], AR1260[554%] due to matrix interference. The MSD {Q1434-02MSD} with File ID: PO109689.D recoveries met the acceptable requirements except for AR1260[348%] due to matrix interference. The Blank Spike met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples K084-1A, K084-2A, K084-2ADL, K084-2B, K084-2BDL, K084-2C, K084-3A, K084-3ADL, K084-3B, K084-3C, K084-4A and K084-4ADL were diluted due to high concentrations.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date