| CHEMTECH PROJECT NUMBER: Q1436 | MATRIX: Solid |
| --- | --- |
| METHOD: 8082A/3541 |  |

|  |  | NA | NO | YES |
| --- | --- | --- | --- | --- |
| 1. | Chromatograms Labeled/Compounds Identified. |  |  | ü |
| 2. | Standard Summary Submitted. |  |  | ü |
| 3. | Calibration - Initial Calibration performed within 30 days before sample analysis and continuing calibration performed within 24 hours of sample analysis, 12 HOURS IF 8000 SERIES METHOD. |  |  |  |
| 4. | Blank Contamination - If yes, list compounds and concentrations in each blank: |  |  |  |
| 5. | Surrogate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable ranges. The Surrogate recoveries met the acceptable criteria except for K084-5AMS [Decachlorobiphenyl(1) - 176%, Decachlorobiphenyl(2) - 205%], K084-5AMSD [Decachlorobiphenyl(1) - 180%, Decachlorobiphenyl(2) - 209%], K084-8A [Decachlorobiphenyl(1) - 8182%, Decachlorobiphenyl(2) - 10482%], K084-8ADL [Decachlorobiphenyl(1) - 12525%, Decachlorobiphenyl(2) - 13435%, Tetrachloro-m-xylene(1) - 154%, Tetrachloro-m-xylene(2) - 146%], K084-8ADL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], K084-9A [Decachlorobiphenyl(1) - 191%, Decachlorobiphenyl(2) - 219%], K084-9ADL [Decachlorobiphenyl(1) - 305%, Decachlorobiphenyl(2) - 333%], K084-10A [Decachlorobiphenyl(1) - 220%, Decachlorobiphenyl(2) - 252%], K084-10ADL [Decachlorobiphenyl(1) - 310%, Decachlorobiphenyl(2) - 352%], K084-10ADL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0%, Tetrachloro-m-xylene(2) - 0%], K084-10B [Decachlorobiphenyl(1) - 212%, Decachlorobiphenyl(2) - 199%], K084-10BDL [Decachlorobiphenyl(1) - 240%, Decachlorobiphenyl(2) - 240%], K084-10BDL2 [Decachlorobiphenyl(1) - 282%, Decachlorobiphenyl(2) - 301%, Tetrachloro-m-xylene(1) - 162%, Tetrachloro-m-xylene(2) - 173%], K084-11A [Decachlorobiphenyl(1) - 316%, Decachlorobiphenyl(2) - 354%], K084-11ADL [Decachlorobiphenyl(1) - 412%, Decachlorobiphenyl(2) - 459%], K084-11ADL2 [Decachlorobiphenyl(1) - 0%, Decachlorobiphenyl(2) - 0%, Tetrachloro-m-xylene(1) - 0% and Tetrachloro-m-xylene(2) - 0%]. Due to high concentration of compounds, these samples required dilution. Therefore, samplew were reanalyzed with dilution and reported |  |  |  |
| 6. | Matrix Spike/Matrix Spike Duplicate Recoveries Meet Criteria If not met, list those compounds and their recoveries which fall outside the acceptable range. The MS {Q1434-03MS} with File ID: PO109562.D recoveries met the requirements for all compounds except for AR1016[182%], AR1260[554%] due to matrix interference. The MS {Q1434-02MS} with File ID: PO109688.D recoveries met the requirements for all compounds except for AR1260[309%] due to matrix interference. The MSD {Q1434-03MSD} with File ID: PO109563.D recoveries met the acceptable requirements except for AR1016[182%], AR1260[554%] due to matrix interference. The MSD {Q1434-02MSD} with File ID: PO109689.D recoveries met the acceptable requirements except for AR1260[348%] due to matrix interference. The Blank Spike met requirements for all samples . The RPD met criteria . |  |  |  |
| 7. | Retention Time Shift Meet Criteria (if applicable) Comments: |  |  |  |
| 8. | Extraction Holding Time Met If not met, list number of days exceeded for each sample: |  |  |  |
| 9. | Analysis Holding Time Met If not met, list those compounds and their recoveries which fall outside the acceptable range. |  |  |  |

ADDITIONAL COMMENTS:

Samples K084-8A, K084-8ADL, K084-8B, K084-9A, K084-10A, K084-10ADL, K084-10B, K084-10BDL, K084-11A and K084-11ADL were diluted due to high concentrations.

The soil samples  results  are based on a dry weight basis.

_______________________________________				__________________________

QA REVIEW								Date